Ethiopian Food and Drug Authority

Guideline for Post-approval Change of Medical Devices

1st Edition

May 2019

Addis Ababa, Ethiopia

- Introduction

Include one paragraph about FDA

Re-write according to article 21 of 1112/2019

Changes in medical devices may take place from time to time as part of their life-cycle. Any change to a registered medical device is linked to the principles of safety, quality and performance and the ability of the regulatory framework to manage the risk of the medical devices.

Before making any decision whether a changed medical device can continue to be placed in the market, the Authority will determine whether evidence of safety, quality and performance have been appropriately collected and reviewed based on the notification made by the registration holder.

For any anticipated change to a medical device, a manufacturer must consider the impact of the change on the patient, practitioner and/or user of the medical device, and the impact of the change on the specifications of the medical device, and decide whether the change is expected to impact the safety, quality and performance of the medical device.

This guideline provides guidance on the categories of changes, the principles of change categorization, and what should be done by the market authorization holder in relation to each category of change to its registered medical device.

- Scope and application

This guideline applies to all registered medical devices. It sets out points for consideration by the market authorization holder when a registered medical device is in the process of change or modification. The guideline also applies to all variations whether from the applicant’s initiative or requested by the Authority. However, it does not apply to medical devices whose application is still under consideration by EFDA.

The guideline specifies the categories of changes in relation to registered medical devices and the requirements to be met to continue the placement of the medical devices in the Ethiopian market.

- Terms and definitions

For the purpose of this document, the following terms and definitions apply.

  - Accessory means an article with an intended purpose as a medical device and that is intended specifically by its manufacturer to be used together with a medical device to enable that medical device to be used in accordance with its intended purpose as a medical device or augment or extend the capabilities of that medical device in fulfillment of its intended purpose.

  - Authority means Ethiopian food and drug authority

  - Change means a post approval variation to any aspect of a medical device, including but not limited to a change to method and site of manufacture, specifications for the finished product, components or raw material, container and container labeling, and product information.

  - Cautions and precautions mean information which alerts the user to exercise special care necessary for the safe and effective use of the medical device. It may include actions to be taken to avoid effects on patients/users that may not be potentially life-threatening or result in serious injury, but about which the user should be aware. Precautions may also alert the user to adverse effects on the device of use or misuse and the care necessary to avoid such effects.

  - Contraindication means a general description of the disease or condition and the patient population for which the device should not be used for the purpose of diagnosing, treating, curing or mitigating. Contraindications are conditions under which the device should not be used because the risk of use clearly outweighs any reasonably foreseeable benefits.

  - Control mechanism means of verifying or checking that the specifications or outputs of the device meet a standard or predetermined result. They are mechanisms put in place to maintain on-going control or regulate the output of a device.

  - Facility means a site that is substantially involved in the manufacture and/or design and manufacture of a medical device.

  - Indications for use means general description of the disease(s) or condition(s) the medical device will diagnose, treat, prevent or mitigate, including where applicable a description of the patient population for which the medical device is intended. The indications include all the labeled uses of the medical device, for example the condition(s) or disease(s) to be prevented, mitigated, treated  or diagnosed, part of the body or type of tissue applied to or interacted with, frequency of use, physiological purpose and patient population.

  - Indirect contact means in relation to the nature of body contact of medical device, includes devices that contact the blood path at one point and serve as a conduit for entry into the vascular system. E.g. blood transfusion tubes, blood bags, etc.

  - Labeling means written, printed or graphic matter presented by a manufacturer meant to provide information concerning a medical device to the users and others, which may be attached to the medical device itself, on its packaging or as a packaging insert or may be made available by other means, for example by electronic means, when appropriate for the purpose as an additional, or alternative way of transmitting certain information regarding the medical device.

  - Operating principles means the means by which a medical device produces or brings about an intended or appropriate effect.

  - Recall means any action taken by the establishment of the device to remove the device and to notify its owners and users of its defectiveness or potential defectiveness, after becoming aware that the device:-

- May be hazardous to health;

- May fail to conform to any claim made by the manufacturer/Authorized Representative relating to its effectiveness, benefits, performance characteristics or safety; or

- May not meet the requirement of the law.

  - Market Authorization Holder means the manufacturer or the authorized representative who applied for and obtained the registration of the medical device under the Proclamation No. 1112/2019.

  - Proclamation means Food and Medicine Administration Proclamation No. 1112/2019

- General principle add some points from singapor, IMDRF

The general principle for categorizing any change to a registered medical device is linked to the principles of safety, quality and performance; and the ability of the regulatory framework to manage the risk of the medical devices. Before making any decision whether a changed medical device can continue to be placed in the market, the Authority will determine whether evidence of safety, quality and performance have been appropriately collected and reviewed based on the notification made by the market authorization holder.

For any anticipated change to a medical device, a manufacturer must consider the impact of the change on the patient, practitioner and/or user of the medical device, and the impact of the change on the specifications of the medical device, and decide whether the change is expected to impact the safety, quality and performance of the medical device.

- Categories of changes

Change to a registered medical device may be categorized into the following three (3) categories:

- Category 1 changes of medical devices that affect their safety, quality and performance and require new registration of the medical device;

- Category 2 changes are changes that require evaluation and endorsement from the EFDA prior to implementation of the change and before placing in the Ethiopian market; and

- Category 3 changes are changes that may be implemented immediately.

For all categories of changes, prior to submission of change notification to the Authority the registration holder may submit a request for confirmation on change category using the template in Annex A. In cases where the category of change cannot be determined or has been deemed inaccurate by the Authority, the Authority shall determine the correct category of change and advice the registration holder to amend the category of change as deemed appropriate. For categories 2 and 3, change notification as according to Annex E is required.

For category 1 changes, the following types of changes require the registration holders to apply for new registration according to medical device registration guide and Medical Device Regulations directive/act (will prepared)

a)     Change to the intended purpose and/or indication of use of a registered medical device, unless it involves a reduction of indications for use not arising due to medical device safety or performance concerns;

b)   Change to the risk classification of a registered medical device;

c)    Change to software that affects safety, quality and performance of the registered medical device;

d) Addition of variant(s) not considered a permissible variant according to the rules of grouping in medical device guideline and IMDRF Product Grouping;

e)     Change to the type, concentration or drug specifications (DS) of medicinal substance in a medical device that incorporates a medicinal product as an ancillary role shall be refer to medicine registration guideline; and

f)      Addition of medical devices with device proprietary names different from the registered devices, into a device listing. Unless the devices with different proprietary names qualify to be listed together less than one listing based on Medical device registration guidance documents on grouping criteria for medical devices registration.

The guiding principles for identification of category 2 of various types of change to registered medical devices are presented in Table 1.

Table 1: Change notification for category 2

| Types of change | Documents to be submitted** |
| --- | --- |
| 5.5.1 Change in manufacturing facility, process and quality management system (QMS) |
| (a) All changes to manufacturing and/or sterilization facilities with no changes to the manufacturing and/or sterilization processes. Example:- Change of manufacturing site. | i) Revised QMS certificate(s) (if applicable); ii) Medical Device labeling stating changes for each amended section (if applicable); iii) Declaration that there is no change to manufacturing and sterilization process; iv) Sterilization validation report. |
| (b) All changes to manufacturing processes (including changes made to outsourced processes) that result in a change in specifications of a registered medical device. Example: Change in the equipment used for cutting the result in the change in length of sutures. Molding or cutting manufacturing process. | Revised QMS certificate(s) (if applicable); Summary of new manufacturing process; Validation report covering new processes; Pre-clinical studies (if applicable); v) Software validation report (for software); vi) Clinical safety report (for operating principles and design characteristics change) (if applicable); vii).Risk analysis. |
| (c) All changes to sterilization processes (including changes made to outsourced processes). Example: Change in moist heat sterilization parameters, or change in sterilization method from ethylene oxide to gamma radiation, or change from batch release to parametric release. | i) Sterilization technique (certificate); ii) Medical Device labeling stating changes for each amended section (if applicable); iii) Sterilization validation report (including the sterilization protocol, sterilization standards applied, sterility assurance level, sterilization revalidation report); iv) QMS certificate(s). |
| 5.5.2 Changes in design or specifications of a registered medical device |
| (a) All changes to the control Mechanisms, operating principles and/or design characteristics of a registered medical device. Example: Change from a quantitative assay to a qualitative assay. Addition of a footswitch to an X-ray system that previously do not operate via a footswitch mechanism. | i) Revised QMS certificate(s) (if applicable); ii) Pre-clinical studies; iii) Risk analysis; iv) Clinical studies (if applicable); v) Medical Device labeling stating changes for each amended section (if applicable); vi) Software validation report (for software, if applicable); vii) Detailed summary of software changes (for software, if applicable). |
| (b) Changes that only involves a design change that does not affect the safety or performance of the medical device (e.g. changes that improve the medical device ergonomics, aesthetic modification of the medical device). | i) Revised QMS certificate(s) (if applicable); ii) Risk analysis; iii) Usability testing report (if applicable). |

| (c) All changes in specifications (including shelf life and stability) of a registered medical device. | i) Revised QMS certificate(s) (if applicable); ii) Pre-clinical studies (if applicable); iii) Clinical safety report (if applicable); iv) Risk analysis; v) Medical Device labeling stating changes for each amended section (if applicable); vi) Software validation report (for software, if applicable); vii) Detailed summary of software changes (for software, if applicable). |
| --- | --- |
| (d) Change to software that affect safety, quality and performance of the registered device such that the treatment or diagnosis of the patient is altered. Example: Upgrade of software version changes the performance characteristics like specificity or sensitivity of the diagnostic medical device. | i) Revised QMS certificate(s) (if applicable); ii) Risk analysis; iii) Software validation report; iv) Detailed summary of software changes. |
| 5.5.3 Changes to materials in a general medical device |
| (a) All changes to biological materials that involve a change in type, source, processing and/or supplier of cells, tissues and/or derivatives of animal, human, microbial or recombinant origin without a change in the intended purpose of the biological material. Example: Change in source of hyaluronic acid from Streptococcus zooepidemicus to Streptococcus equi. | i) Revised QMS certificate(s) (if applicable); ii) Pre-clinical studies, including biological safety data; iii) Clinical safety report (if applicable); iv) Information of sources/donors; v) Risk analysis; |
| (b) All changes to materials or material formulation (of non-biological origin), including changes to medical device coating or surface modification techniques, that involve materials that make direct/indirect contact with body tissues and fluids, or are absorbed by the body. Example: Replacement of catheter surface coating from PEBA to PEEK. | i) Revised QMS certificate(s) (if applicable); ii) List of materials making direct/ indirect contact with human body; iii) Pre-clinical studies; iv) Clinical safety report (if applicable); v) Risk analysis. |
| Types of change | Documents to be submitted** |
| (c) All changes to materials that are used for shielding in medical devices emitting ionising radiation. Example: Change in shielding material of X-ray system from lead to tungsten. | i) Revised QMS certificate(s) (if applicable); ii) Information on radiation source; iii) Information on materials for shielding of radiation; iv) Radiation safety test/test report; v) Risk analysis. |
| (d) All changes to the radiation source (e.g. radioisotopes). | i) Revised QMS certificate(s) (if applicable); ii) Information on radiation source; iii) Radiation safety test/test report; iv) Risk analysis. |
| 5.5.4 Changes to materials in an in-vitro diagnostic (IVD) medical device |
| (a) All changes to the radiation source (e.g. radioisotopes in radioimmunoassay). | i) Revised QMS certificate(s) (if applicable); ii) Pre-clinical performance evaluation data; iii) Clinical performance evaluation data; iv) Information on source of material; v) Radiation safety test/test report; vi) Risk analysis. |
| 5.5.5 Changes to labeling of medical devices |
| (a) All changes to the labeling of medical devices that involve addition, removal and/or revision of warnings, precautions and/or contraindications. Example: Minor changes to clarify the existing wording of the warnings and precautions for a device may not trigger the need for approval. However, in the case where these changes add or remove a contraindication, or remove a warning or precaution, an endorsement by the MDA is required. | i) Revised QMS certificate(s) (if applicable); ii) Description of the warnings, precautions and/or contraindications; iii) Reasons for the revision of approved indications; iv) Medical Device labelling for new medical device(s) stating changes for each amended section. |
| (b) Labeling changes that— i) modify the approved method of use; OR ii) involve a change from ‘Professional use only’ to ‘home use’. | i) Revised QMS certificate(s) (if applicable); ii) Preclinical Studies (if applicable); iii) Clinical safety report (if applicable); iv) Software validation report (for software); v) Risk analysis; vi) Medical Device labeling stating changes for each amended section. |
| 5.5.6 Changes to registered medical devices registration information |
| (a) If within the medical device grouping, the change only— i) involves the addition of new medical devices of the same | i) Justification for addition of medical device(s) to be grouped within the registered medical device group; ii) List of configurations of medical devices; |
| Types of change | Documents to be submitted** |
| design, within the existing range of sizes already registered;OR ii) Involves addition of a new medical device with design change that does not affect the safety or performance of the medical device (e.g. Changes that improve medical device ergonomics, aesthetic modification of the medical device). | iii) Regulatory approval documents from the recognized countries (if applicable); iv) Medical Device information; v) Medical Device labeling stating changes for each amended section; vi) Declaration of conformity; vii) Pre-clinical studies (where applicable); viii) Software validation report (for software, if applicable); ix) Manufacturing information (if applicable). |
| (b) If the change only involves an addition of active, with measuring function or sterile Class A medical device accessories that complement the registered medical device as a system. | i) Declaration by registration holder to state - - the added models are class A sterile; - the name of the medical device affected; - the medical device identifier; - no change in manufacturer for the class A sterile medical device; - name and address for the manufacturing site(s) for class A sterile medical device; ii) List of configurations of medical device; iii) Declaration of conformity; iv) Validation report and certificate. |
| c) All changes to medical device registration that involves an increase or reduction in the number of medical devices in a set grouping of a registered medical device. | i) Declaration of conformity; ii) Declaration from manufacturer to state that there is no change to medical device in all aspects, including intended use, technical specifications; iii) List of configurations of medical devices; iv) Device labeling stating changes for each amended section; v) Description of the addition or reduction. |
| d) All changes to the medical device name and/or medical device identifier. | i) Declaration of conformity; ii) Declaration from manufacturer to state that there is no change to medical device in all aspects, including intended use, technical specifications; iii) List of configurations of medical device; iv) Device labelling stating changes for each amended section. |
| **need to have medical device directive that mention this section, the Authority may, in writing, at any time after the receipt of an application under subsection (1), request the applicant to give to the Authority within the period specified in the request additional information, particulars or document on the application or sample of the medical device; and** need to have medical device directive that mention this section, if any additional information, particulars or document, or sample of the medical device required under subsection (4) is not given by the applicant within the period specified in the request or any extension of time granted by the Authority, the application shall be deemed to be withdrawn and shall not be further proceeded with, but without affecting the right of the applicant to make fresh application. |

5.6    The guiding principles for identification of category 3 of various types of change to registered medical devices are presented in Table 2.

Table 2: Change category for category 3

| Types of change | Documents to be submitted** |
| --- | --- |
| 5.6.1 Change in manufacturing facility, process and quality management system (QMS) |
| (a) All changes to certificates for manufacturing and sterilization facilities that - i) involves an update of certificate QMS validity date only OR; ii) change in scope of the QMS certification which affect the registered medical device (that is not due to safety, and/or performance of the medical device) OR; iii) involves a cancellation of QMS scope on the certificate for any of the multiple existing manufacturing facilities that is related to the registered medical device (that is not due to safety, and/or performance of the medical device), OR; iv) involves the change in conformity assessment body with no change in scope of the certification OR; v) Involves the expansion of scope of the QMS certification which does not affect the registered medical device. | Valid QMS certificate and report. |
| Types of change | Documents to be submitted** |
| 5.6.2 Changes in design or specifications of a registered medical device |
| All changes in software related to design or specifications of a registered medical device are in Category 1, (a) Unless the change only involves a change to software version number that does not affect safety or performance of the medical device, such as— i) software changes solely to correct an inadvertent software error which does not add new functions, does not pose any safety risk and is intended to bring the system to its original specification; ii) software changes which augment interfacing to other non-medical peripherals such as printers or VDUs and which has no diagnostic or therapeutic function; or iii) software changes which only modify the appearance of the user interface with no risk to diagnostic or therapeutic function of the medical device. | i) Software validation report. ii) Detailed summary of software changes. The change notification for this item may be consolidated for a maximum period of 6 months. |
| 5.6.3 Changes to labeling of medical devices |
| (a) Where the change only involves a reduction of indications for use not arising due to medical device safety or performance concerns. | i) Description of the new indications for use; ii) Reasons for the reduction of approved indications; iii) Medical Device labeling for new medical device(s) stating changes for each amended section. |
| (b) Labeling changes that only— i) Involve the addition of Recognized Countries’ approvals (e.g. CE marking). | i) Medical Device labeling stating changes for each amended section; ii) Valid certificates from relevant bodies (where applicable). |
| (c) Other labeling changes involving information in the labeling that does not fall under above (a) and (b). | i) Medical Device labeling stating changes for each amended section; ii) Details of changes and the reason for changes; |
| Types of change | Documents to be submitted** |
| Rephrasing information/ Change in arrangement in IFU/ Change of color/font size/ location of information/ correction of spelling mistake or any administrative change (e.g. from Rd. to road), for example, do not required change notification. Example: Minor changes to clarify the existing wording of the warnings, precautions, and/or how to use for a device in the IFU. | iii) Documents supporting proposed changes detailed above (if applicable). |
| 5.6.4 Changes to registered medical devices registration information |
| (a) If the change only involves an addition of Class A medical device accessories that complement the registered medical device as a system or family. | i) Declaration by registration holder to state - - the added models are class A non-sterile; - the name of the medical device affected; - the medical device identifier; - no change in manufacturer for the class A non-sterile medical device; - name and address for the manufacturing site(s) for class A non-sterile medical device; ii) List of configurations of medical device; iii) Declaration of conformity. |
| b) All deletions of a medical device from medical device registration (for medical devices in grouping). Example: The change only involves the reduction in the number of medical device in the grouping due to obsolescence and not due to safety or performance considerations. | i) Justification for deletion of medical device(s) to be grouped within the registered medical device; ii) Declaration from manufacturer to state that there is no change to medical device in all aspects, including intended use, technical specifications; iii) List of configurations of medical devices; iv) Device labeling stating changes for each amended section. |
| c) All changes in the manufacturer information including changes in manufacturer’s name and address (which is not the manufacturing site). | i) Declaration of conformity; ii) Declaration from manufacturer to state that they will undertake responsibility to provide post market support and assistance related to the medical devices already supplied under the former manufacturer’s name (if applicable); iii) Medical Device labeling stating changes for each amended section. |
| Types of change | Documents to be submitted** |
| d) A change in regulatory status on rejection or withdrawal in any recognized countries for any registered medical device. | i) Existing regulatory approval; ii) Documents from relevant regulatory authorities citing reason for the change in regulatory status; iii) Reason for company to withdraw from regulatory authorities (if applicable). |
| ** need to have medical device directive that mention this section, the Authority may, in writing, at any time after the receipt of an application under subsection (1), request the applicant to give to the Authority within the period specified in the request additional information, particulars or document on the application or sample of the medical device; and ** need to have medical device directive that mention this section, if any additional information, particulars or document, or sample of the medical device required under subsection (4) is not given by the applicant within the period specified in the request or any extension of time granted by the Authority, the application shall be deemed to be withdrawn and shall not be further proceeded with, but without affecting the right of the applicant to make fresh application. |

5.7    Changes to medical devices due to a mandatory reportable incident and/or field corrective action

Changes to medical devices may arise from the occurrence of mandatory reportable incidents and/or field corrective actions (will mention in the medical device directives). The proposed changes to the medical device in these situations may have an impact on the safety and/or performance of the medical device.

The  determination of  the category of  change for  notifications  in  the context  of,  or  as  a consequence of or arising from mandatory reportable incidents and/or field corrective actions shall be based on Clause 5.

6   Notification process

6.1    For Category 2 and 3 changes, each submission of notification for changes is allowed only for one medical device registration.

6.2    Submission of notification for changes shall be made using the form in Annex B.

6.3    All submission of notification of changes shall be accompanied with a fee as per EFDA service directive. Payment shall be made by the applicant prior to endorsement/acknowledgement by EFDA.

7   Requirements for change notification

Refer Annex E for change notification process flow.

a. Registration holder is required to submit completed copies of the following documentation:

i) Request for confirmation on change notification category for registered medical device (Annex A), if necessary. Refer 5.2.

ii)   Change Notification Form (Annex B).

iii)  Summary Table of Change Notification (Annex C).

- Medical Device Safety, quality and performance Declaration (Annex D).

- All supporting documents listed in Annex A, Annex B, Annex C and Annex D.

- Updated Common Submission Dossier Template (CSDT).

- Registration holders are reminded that the determination of documents required for change notification should be made with reference to all submitted changes, and not solely on one category of change.

- EFDA GMP(site Audit)/web for SRA/ certificate for manufacturing site change(for GMP required products such as HIV test kit, malaria, condom )

- Upon the successful submission of the change notification, further amendment of the application will be allowed only Once (within the same change category).

8   Turn around Time

The turn-around time per application for:

a)      Changes in Category 3 is 15 working days. b)      Changes in Category 2 is 30 working days.

Upon submission of complete application form and supporting documents.

ANNEX A:

(Normative)

Request for confirmation on change notification category for registered medical device

[To be printed on Company Letterhead of the registration holder]

Chief Executive

Medical Device Authority

[Your reference number]

Company address                                                                                       [Date]

Dear Sir/Madam,

Request for confirmation on change notification category for registered medical device

The registration holder of the medical device(s) stated below, hereby request for confirmation on change notification category for registered medical device(s) as per description stated below:

| Establishment Licence No. | : |
| --- | --- |
| Medical Device Registration Certificate No. | : |
| MEDCAST Registration Submission ID | : |
| Medical Device Name and include medical device identifier | : |
| Proposed Change Category | :  Category 1  Category 2  Category 3 |
| Description of change | : |
| Relevant Document (attached) | : |

I am aware that a false declaration is an offence under medical device regulation # and may result in the cancellation of registration of the above medical devices under regulation # section ….

I hereby submit a processing fee of Birr . Yours Faithfully,

[Signature]

[Responsible person of Authority establishment] [Company stamp]

Note.

Complete submission for request for confirmation on change notification category will be processed within 14 working days.

ANNEX B

(Normative)

Change Notification Form

Please fill in medical device registration details:

Establishment License No.:

Medical Device Registration Certificate No.: MEDCAST Registration Submission ID:

Medical Device Name and include medical: Device identifier

This checklist is provided to assist the registration holder in determining the required documents for each type of change to be made to any registered medical device.

All the required documents shall be submitted for the relevant sections of the CSDT to support the proposed change to the medical device. The documentation requirements in this checklist are meant to cover the broadest aspect for each category of change. If any required documents as defined in the respective category in this checklist are not available or applicable for the change proposed, please provide a clarification or justification as appropriate.

Please also refer to other relevant documents issued by the Authority which may be required for the registered medical devices.

Table 1. Change notification for category 2

| Types of change | Documents to be submitted** | Provided |
| --- | --- | --- |
| 5.5.1 Change in manufacturing facility, process and quality management system (QMS) | Yes | No (Please Justify) |
| (a) All changes to manufacturing and/or sterilization facilities with no changes to the manufacturing and/or sterilization processes. Example: Change of manufacturing site. | i) Revised QMS certificate(s) (if applicable); ii) Medical Device labeling stating changes for each amended section (if applicable); iii) Declaration that there is no change to manufacturing and sterilization process; iv) Sterilization validation report. |  |  |
| (b) All changes to manufacturing processes (including changes made to outsourced processes) that result in a change in specifications of a registered medical device. Example: Change in the equipment used for cutting the result in the change in length of sutures. Molding or cutting manufacturing process. | i) Revised QMS certificate(s) (if applicable); ii) Summary of new manufacturing process; iii) Validation report covering new processes; iv) Pre-clinical studies (if applicable); v) Software validation report (for software); vi) Clinical safety report (for operating principles and design characteristics change) (if applicable); vii) Risk analysis. |  |  |

| Types of change | Documents to be submitted** | Provided |
| --- | --- | --- |
| (c) All changes to sterilisation processes (including changes made to outsourced processes). Example: Change in moist heat sterilisation parameters, or change in sterilisation method from ethylene oxide to gamma radiation, or change from batch release to parametric release. | i) Sterilisation technique (certificate); ii) Medical Device labelling stating changes for each amended section (if applicable); iii) Sterilisation validation report (including the sterilisation protocol, sterilisation standards applied, sterility assurance level, sterilisation revalidation report); iv) QMS certificate(s); v) Post sterilization functional test report. |  |  |
| 5.5.2 Changes in design or specifications of a registered medical device |
| (a) All changes to the control mechanisms, operating principles and/or design characteristics of a registered medical device Example: Change from a quantitative assay to a qualitative assay Addition of a footswitch to an X-ray system that previously do not operate via a footswitch mechanism. | i)Revised QMS certificate(s) (if applicable); ii) Pre-clinical studies; iii) Risk analysis; iv) Clinical studies (if applicable); v) Medical Device labelling stating changes for each amended section (if applicable); vi) Software validation report (for software, if applicable); vii) Detailed summary of software changes (for software, if applicable). |  |  |
| (b) Changes that only involves a design change that does not affect the safety or performance of the medical device (e.g. changes that improve the medical device ergonomics, aesthetic modification of the medical device). | i) Revised QMS certificate(s) (if applicable); ii) Risk analysis; iii) Usability testing report (if applicable). |  |  |
| (c) All changes in specifications (including shelf life and stability) of a registered medical device | i) Revised QMS certificate(s) (if applicable); ii) Pre-clinical studies (if applicable); iii) Clinical safety report (if applicable); iv) Risk analysis; v) Medical Device labelling stating changes for each amended section (if applicable); vi) Software validation report (for software, if applicable); vii) Detailed summary of software changes (for software, if applicable). |  |  |
| Types of change | Documents to be submitted** | Provided |
| (d) Change to software that affects safety, quality and performance of the registered device such that the treatment or diagnosis of the patient is altered. Example: Upgrade of software version changes the performance characteristics like specificity or sensitivity of the diagnostic medical device. | i) Revised QMS certificate(s) (if applicable); ii) Risk analysis; iii) Software validation report; iv) Detailed summary of software changes. |  |  |
| 5.5.3 Changes to materials in a general medical device |
| (a) All changes to biological materials that involve a change in type, source, processing and/or supplier of cells, tissues and/or derivatives of animal, human, microbial or recombinant origin without a change in the intended purpose of the biological material. Example: Change in source of hyaluronic acid from Streptococcus zooepidemicus to Streptococcus equi. | i) Revised QMS certificate(s)(if applicable); ii) Pre-clinical studies, including biological safety data; iii) Clinical safety report (if applicable); iv) Information of sources/donors; v) Risk analysis. |  |  |
| (b) All changes to materials or material formulation (of non- biological origin), including changes to medical device coating or surface modification techniques, that involve materials that make direct/indirect contact with body tissues and fluids, or are absorbed by the body. Example: Replacement of catheter surface coating from PEBA to PEEK. | i) Revised QMS certificate(s) (if applicable); ii) List of materials making direct/ indirect contact with human body; iii) Pre-clinical studies; iv) Clinical safety report (if applicable); v) Risk analysis. |  |  |
| (c) All changes to materials that are used for shielding in medical devices emitting ionising radiation. Example: Change in shielding material of X-ray system from lead to tungsten. | i) Revised QMS certificate(s) (if applicable); ii) Information on radiation source; iii) Information on materials for shielding of radiation; iv) Radiation safety test/test report; v) Risk analysis. |  |  |
| (d) All changes to the radiation source (e.g. radioisotopes). | i) Revised QMS certificate(s) (if applicable); ii) Information on radiation source; iii) Radiation safety test/test report; iv) Risk analysis. |  |  |

| Types of change | Documents to be submitted** | Provided |
| --- | --- | --- |
| 5.5.4 Changes to materials in an in-vitro diagnostic (IVD) medical device |
| (a) All changes to the radiation source (e.g. radioisotopes in radioimmunoassay). | i) Revised QMS certificate(s) (if applicable); ii) Pre-clinical performance evaluation data; iii) Clinical performance evaluation data; iv) Information on source of material; Radiation safety test/test report; Risk analysis. |  |  |
| 5.5.5 Changes to labeling of medical devices |
| (a) All changes to the labeling of medical devices that involve addition, removal and/or revision of warnings, precautions and/or contraindications. Example: Minor changes to clarify the existing wording of the warnings and precautions for a device may not trigger the need for approval. However, in the case where these changes add or remove a contraindication, or remove a warning or precaution, an endorsement by the MDA is required. | i) Revised QMS certificate(s) (if applicable); ii) Description of the warnings, precautions and/or contraindications; iii) Reasons for the revision of approved indications; iv) Medical Device labeling for new medical device(s) stating changes for each amended section. |  |  |
| (b) Labelling changes that— i) modify the approved method of use; OR ii) involve a change from ‘professional use only’ to ‘home use’. | i) Revised QMS certificate(s) (if applicable); ii) Preclinical Studies (if applicable); iii) Clinical safety report (if applicable); iv) Software validation report (for software); v) Risk analysis; vi) Medical Device labelling stating changes for each amended section. |  |  |
| 5.5.6 Changes to registered medical devices registration information |
| (a) If within the medical device grouping, the change only— i) involves the addition of new medical devices of the same design, within the existing range of sizes already registered; OR ii) Involves addition of a new medical device with design change that does not affect the safety or performance of the medical device (e.g. changes that improve medical device ergonomics, aesthetic modification of the medical device). | Justification for addition of medical device(s) to be grouped within the registered medical device group; ii) List of configurations of medical devices; iii) Regulatory approval documents from the recognized countries (if applicable); iv) Medical Device information; v) Medical Device labeling stating changes for each amended section; vi) Declaration of conformity; vii) Pre-clinical studies (where applicable); viii) Software validation report (for software, if applicable); ix) Manufacturing information (if applicable). |  |  |
| (b) If the change only involves an addition of active, with measuring function or sterile Class A medical device accessories those complement the registered medical device as a system. | i) Declaration by registration holder to state— - the added models are class A sterile; - the name of the medical device affected; - the medical device identifier; - no change in manufacturer for the class A sterile medical device; - name and address for the manufacturing site(s) for class A sterile medical device; ii) List of configurations of medical device; iii) Declaration of conformity; iv) Validation report and certificate. |  |  |
| (c) All changes to medical device registrations that involve an increase or reduction in the number of medical devices in a set grouping of a registered medical device. | i) Declaration of conformity; ii) Declaration from manufacturer to state that there is no change to medical device in all aspects, including intended use, technical specifications; iii) List of configurations of medical devices; iv) Device labelling stating changes for each amended section; v) Description of the addition or reduction. |  |  |
| (d) All changes to the medical device name and/or medical device identifier. | i) Declaration of conformity; ii) Declaration from manufacturer to state that there is no change to medical device in all aspects, including intended use, technical specifications; iii) List of configurations of medical device; iv) Device labeling stating changes for each amended section. |  |  |
| **need to medical device directive section that mention this, the Authority may, in writing, at any time after the receipt of an application under subsection (1), request the applicant to give to the Authority within the period specified in the request additional information, particulars or document on the application or sample of the medical device; and ** need to medical device directive section that mention this, if any additional information, particulars or document, or sample of the medical device required under subsection (4) is not given by the applicant within the period specified in the request or any extension of time granted by the Authority, the application shall be deemed to be withdrawn and shall not be further proceeded with, but without affecting the right of the applicant to make fresh application. |

Table 2: Change category for category 3

| Types of change | Documents to be submitted** | Provided |
| --- | --- | --- |
| 5.6.1 Change in manufacturing facility, process and quality management system (QMS) |
| (a) All changes to certificates for manufacturing and sterilization facilities. | Valid certificate and report. |  |  |
| (b) Unless the change only— i) involves an update of certificate QMS validity date only OR; ii) change in scope of the QMS certification which affect the registered medical device (that is not due to safety, and/or performance of the medical device) OR; iii) involves a cancellation of QMS scope on the certificate for any of the multiple existing manufacturing facilities that is related to the registered medical device (that is not due to safety, and/or performance of the medical device), OR; iv) involves the change in conformity assessment body with no change in scope of the certification, OR; v) Involves the expansion of scope of the QMS certification which does not affect the registered medical device. | Valid QMS certificate. |  |  |
| 5.6.2 Changes in design or specifications of a registered medical device |
| All changes in software related to design or specifications of a registered medical device are in Category 1. (a) Unless the change only involves a change to software version number that does not affect safety or | i) Software validation report; ii) Detailed summary of software changes. The change notification for this item may be consolidated for a maximum period of 6 months. |  |  |
| Types of change | Documents to be submitted** | Provided |
| performance of the medical device, such as— i) software changes solely to correct an inadvertent software error which does not add new functions, does not pose any safety risk and is intended to bring the system to its original specification; ii) software changes which augment interfacing to other non-medical peripherals such as printers or VDUs and which has no diagnostic or therapeutic function; or iii) Software changes which only modify the appearance of the user interface with no risk to diagnostic or therapeutic function of the medical device. |  |  |  |
| 5.6.3 Changes to labelling of medical devices |
| (a) Where the change only involves a reduction of indications for use not arising due to medical device safety or performance concerns. | i) Description of the new indications for use; ii) Reasons for the reduction of approved indications; iii) Medical Device labelling for new medical device(s) stating changes for each amended section. |  |  |
| (b) Labeling changes that only— i) involve the addition of Recognized Countries’ approvals (e.g. CE marking). | i) Medical Device labeling stating changes for each amended section; ii) Valid certificates from relevant bodies (Where applicable). |  |  |
| (c) Other labelling changes involving information in the labelling that does not fall under (a) and (b). Rephrasing information/ Change in arrangement in IFU/ Change of colour/ font size / location of information/ correction of spelling mistake or any administrative change (e.g. | i) Medical Device labelling stating changes for each amended section; ii) Details of changes and the reason for changes; iii) Documents supporting proposed changes detailed above (if applicable). |  |  |

| Types of change | Documents to be submitted** | Provided |
| --- | --- | --- |
| from Rd. to road), for example, do not required change notification. Example: Minor changes to clarify the existing wording of the warnings, precautions, and/or how to use for a device in the IFU. |  |  |  |
| 5.6.4 Changes to registered medical devices registration information |
| (a) If the change only involves an addition of Class A medical device accessories that complement the registered medical device as a system or family. | i) Declaration by registration holder to state - - the added models are class A non- sterile; - the name of the medical device affected; - the medical device identifier; - no change in manufacturer for the class A non-sterile medical device; - name and address for the manufacturing site(s) for class A non- sterile medical device; ii) List of configurations of medical device; iii) Declaration of conformity. |  |  |
| (b) All deletions of a medical device from medical device registration (for medical devices in grouping). Example: The change only involves the reduction in the number of medical device in the grouping due to obsolescence and not due to safety or performance considerations. | i) Justification for deletion of medical device(s) to be grouped within the registered medical device; ii) Declaration from manufacturer to state that there is no change to medical device in all aspects, including intended use, technical specifications; iii) List of configurations of medical devices; iv) Device labeling stating changes for each amended section. |  |  |
| (c) All changes in the manufacturer information including changes in Manufacturer’s name and address (which is not the manufacturing site). | i) Declaration of conformity; ii) Declaration from manufacturer to state that they will undertake responsibility to provide post market support and assistance related to the medical devices already supplied under the former manufacturer’s name (if applicable); iii) Medical Device labeling stating changes for each amended section. |  |  |
| (d) A change in regulatory status on rejection or withdrawal in any recognized countries for any registered medical device. | i) Existing regulatory approval; ii) Documents from relevant regulatory authorities citing reason for the change in regulatory status; iii) Reason for company to withdraw from regulatory authorities (if applicable). |  |  |
| ** need to medical device directive section that mention this, the Authority may, in writing, at any time after the receipt of an application under subsection (1), request the applicant to give to the Authority within the period specified in the request additional information, particulars or document on the application or sample of the medical device; and ** need to medical device directive section that mention this, if any additional information, particulars or document, or sample of the medical device required under subsection (4) is not given by the applicant within the period specified in the request or any extension of time granted by the Authority, the application shall be deemed to be withdrawn and shall not be further proceeded with, but without affecting the right of the applicant to make fresh application. |

ANNEX C

(Informative)

Summary Table of Changes

This annex provides guidelines on completing the Summary Table of Change Notification.

(a) This summary table is to be completed and submitted for all change applications.

(b) List the proposed changes, according to the “category of change”, to the registered medical device(s) in the summary table below. All applicable types of changes are to be included; any change not specified in this table will not be included for the change notification.

(c) Information to be included in the table is explained below:

i)   Type of changes: Please state clearly the type of change, category of change and medical device registration number certificate #.

-   With reference to the ‘type of changes’ categories in Table 1, highlight the type of change proposed.

-   Specify the IRIS@ medical device registration number/certificate number for the registered medical device(s) included in this change (if the proposed change is identical and applicable to identical medical devices across multiple device registrations on the IRIS@; list the applicable medical device registrations). Confirm these medical device(s) subjected to the change.

NOTE All applicable types of changes are to be included. If the types of change proposed affects/results in another type of change, all types of changes shall be included. For

Example, change in material of medical device and change (update) of labeling often occur together.

ii)   Present: Please state clearly the current scope and aspects of the medical device to be changed.

iii)  Proposed: Please state clearly the proposed scope and aspects of change.

iv)  Reason for change: Please state clearly the rationale for the proposed scope and aspects of change.

v)  Status of proposed change in recognized countries: Please state the reference agency status (approved/authorized for marketing) for these proposed changes.

24

| (a) Type of changes | (b) Present | (c) Proposed | (d) Reason for change | (e) Status of proposed change in recognised countries |
| --- | --- | --- | --- | --- |
| Type of change: e.g. Change in material: Delivery tube material changed from polyvinyl chloride(PVC) to silicone Category of change: | Delivery tube material: polyvinylchloride (PVC) Registration no: List of medical device i) ii) iii) | Delivery tube material: silicone | Improve patient safety by changing to DEHP-free tubing material | Australia TGA –pending EU Notified Body – approved/authorised for marketing Health Canada – not supplied US FDA – not supplied Japan MHLW – not supplied |
| Type of change: e.g. Change in manufacturing facility Category of change: | Name and address of current manufacturing facility A Registration no: List of medical device i) ii) iii) | Name and address of new manufacturing facility B | Reason for to move manufacturing activities from facility A to facility B | Australia TGA – pending EU Notified Body – approved/authorised for marketing Health Canada – not supplied US FDA – not supplied Japan MHLW – not supplied |

ANNEX D

(Normative)

Medical Device Safety, quality and performance Declaration Template

[To be printed on Company Letterhead of the registration holder]

EFMHACA

[Date]

Dear Sir/Madam,

Declaration of Medical Device Safety, quality and performance on Change Notification

I, on behalf of [company name], the manufacturer of the medical device(s) stated below, hereby declare that the medical device(s) in this change notification, is/are not a subject of a mandatory reportable incident and/or an ongoing field corrective action conform(s) to the Essential Principles for Safety, quality and performance as per the statutory requirements of the Medical Device GUIDELINE and the Medical Device Regulations Directive (will prepared).

This declaration shall apply to the following medical device(s):

[List containing medical devices names and registration submission ID]

I am aware that a false declaration is an offence under proclamation 1112/19(_will prepared) and may result in the cancellation of registration of the above medical devices under Section of proclamation/directive.

Yours Faithfully,

[Signature]

[Full Name and Title (Top Management Official)]                         [Company stamp]

NB.:-

Applicant pick the respective change notification category (refer to Section 7 of guidance document)

- Category 1 changes of medical devices that affect their safety, quality and performance and require new registration of the medical device;

- Category 2 changes are changes that require evaluation and endorsement form EFMHACA prior to implementation of the change and before placing in the market

- Category 3 changes may be implemented immediately.

Reference:- IMDRF and Malaysian change notification guide