Ethiopian Food and Drugs Authority

Common Submission Dossier Template for In Vitro Diagnostic Medical Devices

May 2019

Addis Ababa, Ethiopia

TABLE OF CONTENTS

PREFACE	3

INTRODUCTION	3

SCOPE	3

PURPOSE	3

CLASSIFICATION MATRICES	3

DEFINITIONS	3

NUMBERING OF HEADINGS	4

QUALITY MANAGEMENT SYSTEM CHAPTERS (6A & 6B)	4

LANGUAGE REQUIREMENTS	4

PAGINATION	4

OTHER GENERAL NOTES	5

ACRONYMS	6

HIERARCHY PRESENTATION	7

CHAPTER 1 – REGIONAL ADMINISTRATIVE	11

CHAPTER 2 – SUBMISSION CONTEXT	19

CHAPTER 3 – ANALYTICAL PERFORMANCE AND OTHER EVIDENCE	24

CHAPTER 4 – CLINICAL EVIDENCE	41

CHAPTER 5 – LABELLING AND PROMOTIONAL MATERIAL	44

CHAPTER 6A – QUALITY MANAGEMENT SYSTEM PROCEDURES	47

CHAPTER 6B – QUALITY MANAGEMENT SYSTEM DEVICE SPECIFIC INFORMATION	49

DOCUMENT REVISION HISTORY	52

INTRODUCTION

So far, the Medical Devices general registration requirements have been part of the Registration Guideline, which was adopted since 2014. EFDA’s Medical Devices Registration and Licensing team has now established a comprehensive harmonized structure for premarket medical device submissions that assists the manufacturers to submit uniform registration dossier when registering their products.

This document provides harmonized, modular, format for use when filing IVD medical device submissions to EFDA for market authorization. Irrespective of the requirements of this Guidance Document, EFDA has the right to request for information or material, or define conditions not specifically described in this document that is deemed necessary for the purpose of regulatory control.

Manufacturers of all classes of medical device are expected to demonstrate conformity of the device to the EPSP through collection and examination of evidence of conformity in technical documentation that shows how each medical device was developed, designed and manufactured together with the descriptions and explanations necessary to understand the manufacturer’s determination with respect to such conformity. This technical documentation must be updated as necessary to reflect the current status, specification and configuration of the device.

The evidence of conformity of medical device to the EPSP must be compiled in the format of CSDT for the purpose of conformity assessment and submission of application for medical device registration to provide evidence to the Authority or Conformity Assessment Body (CAB) that the subject medical device is in conformity with the Essential Principles. The CSDT need to be kept in the premise for audit or inspection purposes.

EFDA’s medical devices registration and licensing team will monitor the use of these structures and work to continually improve the documents at appropriate intervals based on sufficient use and experience. Comments or questions associated with these documents will be accepted via regulatory@fmhaca.gov.et.

SCOPE

This document applies for In-vitro diagnostics device (IVD) market authorization submissions.

PURPOSE

To create a comprehensive submission structure that can be used as a harmonized electronic/softcopy submission format while minimizing divergences from manufacturers to manufacturers when submitting their registration dossiers. This document is intended to provide guidance regarding the location of submission elements.

DEFINITIONS

FULL REPORT - Typically includes a complete, detailed description of the objective of the assessment, the methods and procedures, study endpoint(s), pre-defined pass/fail criteria, deviations, results summary, discussion and conclusions, and may include data. Complete, detailed support of method selection, worst case justification, study endpoint selection, and pass/fail criteria should be included.

SUMMARY - Typically includes a brief synopsis of the assessment (1) purpose, (2) methods and (3) results and (4) discussion and conclusions. Outliers and deviations should be reported with the results. The purpose of the assessment and description of methods should address:

- Why the characteristic being evaluated is of interest; and

- Why the particular methods are being used to evaluate the characteristic, if applicable including why a harmonized/recognized standard has or has not been complied with.

SUBMISSION – A regulatory submission can be any type of information related to a medical device regulatory process. This includes but is not limited to a request for approval/authorization to market a device, any communications relating to the original submission, and any request for modification to an existing approval.

QUALITY MANAGEMENT SYSTEM CHAPTERS (6A & 6B)

Chapter 6A & B of the ToC is written in terms of the quality management system language employed in ISO 13485. Chapter 6A is where the company places the standard operating procedures (SOPs) the company utilizes to implement its overall high level quality management system. Chapter 6B is where the company places the documents and records the company utilizes to implement the quality management system SOPs described in Chapter 6A.

LANGUAGE REQUIREMENTS

The full dossier files must be written in English or Amharic.

PAGINATION

Pages of the submission should be numbered in such a manner that information can be referenced by page number. This may be done either by consecutively numbering the entire submission, or numbering the pages within a section or chapter (e.g., CH2.4.1-1, CH2.4.1-2).

OTHER GENERAL NOTES

This outline of documentation is to support a smooth documentation process. It remains the applicant’s responsibility to ensure all regulatory requirements are met, and that clear and transparent evidence of conformity to these requirements are provided.

The whole submission document shall be as a single PDF file. The whole contents shall be in text format except scanned certificates and pictures.

ACRONYMS

| EFDA | Ethiopian Food and Drugs Administration |
| --- | --- |
| CAPA | Corrective Action and Preventive Action |
| CSDT | Common Submission Dossier Template |
| EU | European Union |
| GMDN | Global Medical Device Nomenclature |
| IMDRF | International Medical Device Regulators Forum |
| MDUFA | Medical Device User Fee Amendments |
| NB | Notified Body |
| SUD | Single Use Device |
| TGA | Therapeutic Goods Administration - Australia |
| USFDA | United States Food and Drug Administration |
| IRB | Institutional Review Board |

HIERARCHY PRESENTATION

The following is a hierarchical presentation of the submission structure. More detailed guidance regarding where elements belong is provided following this table.

CHAPTER 1 – REGIONAL ADMINISTRATIVE

| CH1.01 | Cover Letter |
| --- | --- |
| CH1.02 | Submission Table of Contents |
| CH1.03 | List of Terms/Acronyms |
| CH1.04 | Application Form/Administrative Information |
| CH1.05 | Listing of Device(s) |
| CH1.06 | Quality Management System, Full Quality System or other Regulatory Certificates |
| CH1.07 | Free Sale Certificate/ Certificate of Marketing authorization |
| CH1.08 | User Fees |
| CH1.09 | Pre-Submission Correspondence and Previous Regulator Interactions |
| CH1.10 | Acceptance for Review Checklist |
| CH1.11 | Statements/Certifications/Declarations of Conformity |
| CH1.11.1 |  | Performance and Voluntary Standard |
| CH1.11.2 |  | Environmental Assessment |
| CH1.11.3 |  | Clinical Trial Certifications |
| CH1.11.4 |  | Indications for Use Statement with Rx and/or OTC designation Enclosure |
| CH1.11.5 |  | Truthful and Accurate Statement |
| CH1.11.6 |  | Declaration of Conformity |
| CH1.12 | Letters of Reference for Master Files |
| CH1.13 | Letter of Authorization |
| CH1.14 | Other Regional Administrative Information |
| CHAPTER 2 – SUBMISSION CONTEXT |  |
| CH2.1 | Chapter Table of Contents |
| CH2.2 | General Summary of Submission |
| CH2.3 | Summary and Certifications for Premarket Submissions |
| CH2.4 | Device Description |
| CH2.4.1 |  | Comprehensive Device Description and Principle of Operation |
| CH2.4.2 |  | Material Specifications |
| CH2.4.3 |  | Description of Device Packaging |
| CH2.4.4 |  | History of Development |
| CH2.4.5 |  | Reference and Comparison to Similar and/or Previous Generations of the Device |
| CH2.4.6 |  | Substantial Equivalence Discussion |
| CH2.5 | Indications for Use and/or Intended Use |
| CH2.5.1 |  | Intended Use; Intended Purpose; Intended User; Indications for Use |
| CH2.5.2 |  | Intended Environment/Setting for use |
| CH2.5.3 |  | Pediatric Use |
| CH2.5.4 |  | Contraindications for Use |
| CH2.6 | Global Market History |
| CH2.6.1 |  | Global Market History |
| CH2.6.2 |  | Global Incident Reports and Recalls |
| CH2.6.3 |  | Sales, Incident and Recall Rates |
| CH2.6.4 |  | Evaluation/Inspection Reports |
| CH2.7 | Other Submission Context Information |
| CHAPTER 3 – NON-CLINICAL EVIDENCE |  |
| CH3.1 |  | Chapter Table of Contents |
| CH3.2 |  | Risk Management |
| CH3.3 |  | Essential Principles (EP) Checklist |
| CH3.4 |  | Standards |
| CH3.4.1 |  | List of Standards |
| CH3.4.2 |  | Declaration and/or Certification of Conformity |
| CH3.5 |  | Analytical Performance |
| CH3.5.01 |  | Stability of Sample(s) |
| CH3.5.01.1 |  | [Study description, study identifier, date of initiation, date of completion] |
| CH3.5.01.1.1 |  | Summary |
| CH3.5.01.1.2 |  | Full Report |
| CH3.5.01.1.3 |  | Statistical Data |
| CH3.5.02 |  | Validation of Specimens |
| CH3.5.02.1 |  | [Study description, study identifier, date of initiation, date of completion] |
| CH3.5.02.1.1 |  | Summary |
| CH3.5.02.1.2 |  | Full Report |
| CH3.5.02.1.3 |  | Statistical Data |
| CH3.5.03 |  | Metrological traceability of calibrator and control material values |

_______________________________________________________________________________________________________________

| CH3.5.03.1 | [Study description, study identifier, date of initiation, date of completion] |
| --- | --- |
| CH3.5.03.1.1 | Summary |
| CH3.5.03.1.2 | Full Report |
| CH3.5.03.1.3 | Statistical Data |
| CH3.5.04 | Accuracy of Measurement |
| CH3.5.04.1 | Trueness |
| CH3.5.04.1.1 | [Study description, study identifier, date of initiation, date of completion] |
| CH3.5.04.1.1.1 | Summary |
| CH3.5.04.1.1.2 | Full Report |
| CH3.5.04.1.1.3 | Statistical Data |
| CH3.5.04.2 | Precision (Repeatability and Reproducibility) |
| CH3.5.04.2.1 | [Study description, study identifier, date of initiation, date of completion] |
| CH3.5.04.2.1.1 | Summary |
| CH3.5.04.2.1.2 | Full Report |
| CH3.5.04.2.1.3 | Statistical Data |
| CH3.5.05 | Analytical Sensitivity |
| CH3.5.05.1 | [Study description, study identifier, date of initiation, date of completion] |
| CH3.5.05.1.1 | Summary |
| CH3.5.05.1.2 | Full Report |
| CH3.5.05.1.3 | Statistical Data |
| CH3.5.06 | Analytic Specificity |
| CH3.5.06.1 | [Study description, study identifier, date of initiation, date of completion] |
| CH3.5.06.1.1 | Summary |
| CH3.5.06.1.2 | Full Report |
| CH3.5.06.1.3 | Statistical Data |
| CH3.5.07 | High Dose Hook Effect |
| CH3.5.07.1 | [Study description, study identifier, date of initiation, date of completion] |
| CH3.5.07.1.1 | Summary |
| CH3.5.07.1.2 | Full Report |
| CH3.5.07.1.3 | Statistical Data |
| CH3.5.08 | Measuring Range of the Assay |
| CH3.5.08.1 | [Study description, study identifier, date of initiation, date of completion] |
| CH3.5.08.1.1 | Summary |
| CH3.5.08.1.2 | Full Report |
| CH3.5.08.1.3 | Statistical Data |
| CH3.5.09 | Validation of Assay Cut-off |
| CH3.5.09.1 | [Study description, study identifier, date of initiation, date of completion] |
| CH3.5.09.1.1 | Summary |
| CH3.5.09.1.2 | Full Report |
| CH3.5.09.1.3 | Statistical Data |
| CH3.5.10 | Validation of the Assay Procedure |
| CH3.5.10.1 | [Study description, study identifier, date of initiation, date of completion] |
| CH3.5.10.1.1 | Summary |
| CH3.5.10.1.2 | Full Report |
| CH3.5.10.1.3 | Statistical Data |
| CH3.6 | Other Studies |
| CH3.6.1 | Electrical Systems: Safety, Mechanical and Environmental Protection, and Electromagnetic Compatibility |
| CH3.6.1.1 | [Study description, study identifier, date of initiation, date of completion] |
| CH3.6.1.1.1 | Summary |
| CH3.6.1.1.2 | Full Report |
| CH3.6.1.1.3 | Statistical Data |
| CH3.6.2 | Software/Firmware |
| CH3.6.2.01 | Software/Firmware Description |
| CH3.6.2.02 | Hazard Analysis |
| CH3.6.2.03 | Software Requirement Specification |
| CH3.6.2.04 | Architecture Design Chart |
| CH3.6.2.05 | Software Design Specification |
| CH3.6.2.06 | Traceability Analysis |
| CH3.6.2.07 | Software Life Cycle Process Description |
| CH3.6.2.08 | Software Verification and Validation |
| CH3.6.2.08.1 | [Study description, study identifier, date of initiation] |
| CH3.6.2.08.1.1 | Summary |
| CH3.6.2.08.1.2 | Full Report |
| CH3.6.2.08.1.3 | Statistical Data |
| CH3.6.2.09 | Revision Level History |
| CH3.6.2.10 | Unresolved Anomalies (Bugs or Defects) |
| CH3.6.2.11 | Cybersecurity |
| CH3.6.2.12 | Interoperability |

_______________________________________________________________________________________________________________

| CH3.6.3 |  | Cleaning and Disinfection Validation |
| --- | --- | --- |
| CH3.6.3.1 |  | [Study description, study identifier, date of initiation, date of completion] |
| CH3.6.3.1.1 |  | Summary |
| CH3.6.3.1.2 |  | Full Report |
| CH3.6.3.1.3 |  | Statistical Data |
| CH3.6.4 |  | Usability/Human Factors |
| CH3.6.4.1 |  | [Study description, study identifier, date of initiation, date of completion] |
| CH3.6.4.1.1 |  | Summary |
| CH3.6.4.1.2 |  | Full Report |
| CH3.6.4.1.3 |  | Statistical Data |
| CH3.6.5 |  | Stability of the IVD |
| CH3.6.5.1 |  | Claimed Shelf-life |
| CH3.6.5.1.1 |  | [Study description, study identifier, date of initiation, date of completion] |
| CH3.6.5.1.1.1 |  | Summary |
| CH3.6.5.1.1.2 |  | Full Report |
| CH3.6.5.1.1.3 |  | Statistical Data |
| CH3.6.5.2 |  | In Use Stability |
| CH3.6.5.2.1 |  | [Study description, study identifier, date of initiation, date of completion] |
| CH3.6.5.2.1.1 |  | Summary |
| CH3.6.5.2.1.2 |  | Full Report |
| CH3.6.5.2.1.3 |  | Statistical Data |
| CH3.6.5.3 |  | Shipping Stability |
| CH3.6.5.3.1 |  | [Study description, study identifier, date of initiation, date of completion] |
| CH3.6.5.3.1.1 |  | Summary |
| CH3.6.5.3.1.2 |  | Full Report |
| CH3.6.5.3.1.3 |  | Statistical Data |
| CH3.7 |  | Analytical Performance and Other Evidence Bibliography |
| CH3.8 |  | Other Evidence |
| CH3.8.1 |  | [Study description, study identifier, date of initiation, date of completion] |
| CH3.8.1.1 |  | Summary |
| CH3.8.1.2 |  | Full Report |
| CH3.8.1.3 |  | Statistical Data |
| CHAPTER 4 – CLINICAL EVIDENCE |
| CH4.1 | Chapter Table of Contents |
| CH4.2 | Overall Clinical Evidence Summary |
| CH4.2.1 |  | Expected Values/Reference Ranges |
| CH4.2.2 |  | Clinical Evidence Evaluation Report |
| CH4.2.3 |  | Device Specific Clinical Studies |
| CH4.2.3.1 |  | [Study description, protocol #, date of initiation, date of completion] |
| CH4.2.3.1.1 |  | Clinical Study Synopsis |
| CH4.2.3.1.2 |  | Clinical Study Report |
| CH4.2.2.1.3 |  | Clinical Study Data |
| CH4.2.4 |  | Clinical Literature Review and Other Reasonable Known Information |
| CH4.3 | IRB Approved Informed Consent Forms |
| CH4.4 | Investigators Sites and IRB contact information |
| CH4.5 | Other Clinical Evidence |
| CH4.5.1 |  | [Study description, study identifier, date of initiation, date of completion] |
| CH4.5.1.1 |  | Summary |
| CH4.5.1.2 |  | Full Report |
| CH4.5.1.3 |  | Statistical Data |
| CHAPTER 5 – LABELLING AND PROMOTIONAL MATERIAL |
| CH5.1 | Chapter Table of Contents |
| CH5.2 | Product/Package Labels |
| CH5.3 | Package Insert/Instructions for Use |
| CH5.4 | e-labelling |
| CH5.5 | Patient Labelling |
| CH5.6 | Technical/Operators Manual |
| CH5.7 | Product Brochures |
| CH5.8 | Other Labelling and Promotional Material |
| CHAPTER 6A – QUALITY MANAGEMENT SYSTEM PROCEDURES |
| CH6A.1 | Cover Letter |
| CH6A.2 | Chapter Table of Contents |
| CH6A.3 | Administrative |
| CH6A.3.1 |  | Product Descriptive Information |
| CH6A.3.2 |  | General Manufacturing Information |
| CH6A.3.3 |  | Required Forms |

_______________________________________________________________________________________________________________

| CH6A.4 | Quality management system procedures |
| --- | --- |
| CH6A.5 | Management responsibilities procedures |
| CH6A.6 | Resource management procedures |
| CH6A.7 | Product realization procedures |
| CH6A.7.1 | Design and development procedures |
| CH6A.7.2 | Purchasing procedures |
| CH6A.7.3 | Production and service controls procedures |
| CH6A.7.4 | Control of monitoring and measuring devices procedures |
| CH6A.8 | QMS measurement, analysis and improvement procedures |
| CH6A.9 | Other Quality System Procedures Information |
| CHAPTER 6B – QUALITY MANAGEMENT SYSTEM DEVICE SPECIFIC INFORMATION |
| CH6B.1 | Chapter Table of Contents |
| CH6B.2 | Quality management system information |
| CH6B.3 | Management responsibilities information |
| CH6B.4 | Resource management information |
| CH6B.5 | Device Specific Quality Plan |
| CH6B.6 | Product realization information |
| CH6B.6.1 | Design and development information |
| CH6B.6.2 | Purchasing information |
| CH6B.6.3 | Production and service controls information |
| CH6B.6.4 | Control of monitoring and measuring devices information |
| CH6B.7 | QMS measurement, analysis and improvement information |
| CH6B.8 | Other Device Specific Quality Management System Information |

CHAPTER 1 – ADMINISTRATIVE

1.1. Cover Letter

a) The cover letter should state applicant or sponsor name and/or their authorized representative, the type of submission, the common name of the device (if applicable), device trade name or proprietary name (both of the base device and a new name if one is given to the new version/model of the device) and include the purpose of the application, including any changes being made to existing approvals.

b) If applicable, it should include information pertaining to any Master Files referenced by the submission.

c) If applicable, acknowledgement that a device sample has been submitted or offered alternatives to allow the Authority to view or access the device (when the Authority requests a sample).

d) If the submission is requesting approval of a change that is the result of CAPA due to a recall, this should be stated.

e) If the submission is in response to a request for information from the Authority, this should be stated and the date of that letter should be included as well as any reference number(s).

f)  If the submission is unsolicited information (where accepted), this should be stated and any related reference number(s) provided.

NOTE: The cover letter should not contain any detailed scientific information.

1.2. Submission Table of contents

- Includes at least level 1 & 2 headings for the entire submission Contents

- Specifies the page number for each item referred to in the table.

NOTE: Refer to the Pagination Section of this document for information about submission pagination.

1.3. Terms/Acronyms

List of Terms or acronyms used in the submission that require definition, should be defined here.

1.4. Application Form/Administrative Information

The application form shall be filled online on (www……). The copy of the filled application form must be included in this submission.

1.5. Listing of Device(s)

A table listing each variant/model/configuration/component/accessory that is the subject of the submission and the following information for each variant/model:

a)  The identifier (e.g. bar code, catalogue, model or part number, UDI)

b)  A statement of its name/description that provides (e.g. Trade name, size, material)

NOTE:

i.   A model/variant/configuration/component/accessory of a device has common specifications, performance and composition, within limits set by the applicant.

ii.    Typically each item listed should be available for sale. For example, if everything is sold as part of a kit, then this list would only include the kit. You do not need to list all components that may be sold within a kit/set, unless the component is available for sale independently of the kit.

iii. The listing should comply with the grouping criteria where the listed devices are allowed to be applied in one application dossier. Refer to Doc …..for grouping principles (hyperlink to the actual document).

1.6. Quality Management System, Full Quality System or Other Regulatory Certificates

The Manufacturer shall have certificates confirming the implementation of good quality management system in the device’s production process.

One of these certificates to be considered is Good Manufacturing Practice Certificate (GMPC) issued by EFDA, covering the scope of products.

EN ISO 13485 certificate in case it is issued by another Notified Body or registrar. CE full quality system certificates covering the scope of products when issued by another Notified Body.

1.7. Free Sale Certificate/Certificate of Marketing Authorization

Document/certificate issued by the Regulatory Authority where the medical device is marketable, attesting that the device is marketable, without any restriction at their jurisdiction.

This certificate shall indicate the name(s) of the device(s) (with model if applicable) and explains whether the products are freely sold in the country of origin should be provided; if not, the reasons thereof should be clearly stated with appropriate justification.

If the manufacturer of the medical device has subsidiaries, a free sale certificate should indicate the name and address of the subsidiaries with the name of the device they manufacture, and/or a separate free sale certificate should be submitted for each subsidiary.

The certificate should be original, valid, and authenticated by the Ethiopian embassy. If the certificate comes from a country where there is no Ethiopian embassy, the Authority will contact directly the responsible body that provides the free sale certificate.

1.8. Service Fees

The applicant shall attach the application (evaluation) fee payment proof to this submission.

1.09. Declarations of Conformity

As part of the conformity assessment procedures, the manufacturer of a medical device is required to make a Declaration of Conformity that declares that the device complies with:

a) The applicable provisions of the Essential Principles/Requirements

b) The classification rules

c) An appropriate conformity assessment procedure

The declaration must include the following information:

Date from which the Declaration of Conformity is valid;

Name and address of the device manufacturer; and,

The name, position, and signature of the responsible person who has been authorized to complete the Declaration of Conformity on behalf of the manufacturer.

1.10. Letter of Reference for Master Files

Letter from any Master File owner granting access to the information in the master file. The letter should specify the scope of access granted.

1.11. Agency Agreement

An Agency Agreement made by and between the device manufacturer/Authorized representative and importer/local agent to market the device in Ethiopia.

CHAPTER TWO

SUBMISSION CONTEXT

2.1. Chapter Table of Contents

a)  Includes all headings and sub-headings for the chapter.

b) Specifies the page number for each item referred to in the table.

2.2. General Summary of Submission

a, Statement of the device type (e.g. hip implant, infusion pump, standalone software) and name (e.g. trade name, proprietary name), its general purpose, and a high-level summary of key supporting evidence (i.e. studies that are unique to the risks of this device type, for example burst testing of a ceramic femoral head; electrical safety evaluation (IEC 60601) testing for an infusion pump).

b) Summary of submission, including

i. The type of submission (e.g. new, amendment, change of existing application, renewal);

ii. If amendment/supplement, the reason of the amendment/supplement;

iii. If a change to existing approval, description of the change requested (e.g., changes in design, performance, indications, changes to manufacturing processes, manufacturing facilities, suppliers);

iv. Any high-level background information or unusual details that the manufacturer wishes to highlight in relation to the device, its history or relation to other approved devices or previous submissions (provides context to submission).

2.4. Device Description	NO CONTENT AT THIS LEVEL

2.4.1 Comprehensive Device Description and Principle of Operation

a)   A general description of the device, including:

- A statement of the device name.

- What does it detect?

- Who uses it and for what? (high level statement)

- Where to use it? (Places/environment where the device is intended to be used)\

- General description of the principle of the assay method or instrument principles of operation

- Description of the components (e.g. reagents, assay controls and calibrators) and where appropriate, a description of the reactive ingredients of relevant components (such as antibodies, antigens, nucleic acid primers).

- If applicable, labelled pictorial representation (diagrams, photos, drawings).

- If system, how the components relate?

- If applicable, identify if the device incorporates software/firmware and its role.

b)   Product specification, including:

i. Physical characteristics of relevance to the end user (dimensions, weight)

ii. If applicable, technical features and operating modes

iii. If applicable, operating specifications and performance characteristics (e.g. electrical power requirements, settings and associated allowable ranges/limits, temperature and humidity limits, number of tests per hour, sensitivity/specificity)

iv. If applicable, a complete list of the configurations/models of the devices and a summary of the differences in specifications (comparison table and/or pictures/diagrams with supporting text).

c)   If applicable, engineering diagrams/prints/schematics of the device.

d)   Describe the different specimen types that can be used for this device (e.g. serum, plasma, urine, cerebrospinal fluid), including any additives that are required (e.g. anticoagulant).

e)   Describe the use of controls. If applicable, a list of compatible control materials or control material specifications.

f)   Description of the accessories, other IVD or non-IVD medical devices and other products, which are intended to be used in combination with the IVD medical device.

g)   If approved by the regulator, provide the approval number and identification for each of the accessories, other IVD or non-IVD medical devices and other products, which are intended to be used in combination with the IVD medical device.

h)   If applicable, indication of biological material or derivate used in the medical device, including: origin (human, animal, recombinant or fermentation products or any other biological material) and source (e.g. blood, bone, heart, any other tissue or cells). Where a significant risk is identified, a brief summary of evaluations performed to minimize biological risks, in particular, with regard to viruses and other transmissible agents.

i)    If the device contains an active pharmaceutical ingredient (API) or drug, an indication of the substance, should be provided.  This should include its identity and source, and the intended reason for its presence and its primary mode of action.

j)    Description of the collection and/or transport container(s) provided with the IVD medical device or a description of specifications or recommended collection and/or transport container(s).

k)   If applicable, a listing of assays that are compatible with the instrument. l)    If applicable, a listing of compatible instruments.

m)  A list of any software to be used with the IVD medical device and a description of its role in the delivery of the intended purpose

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the comprehensive device description and principles of operations provided in this section regarding the subject device.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the comprehensive device description and principles of operations provided in this section regarding the subject device.

2.4.2. Description of Device Packaging

a) Information regarding the packaging of the devices, including, when applicable, primary packaging, secondary and any other packaging associated;

b) Specific packaging of accessories marketed together with the medical devices shall also be described;

c) If the user needs to package the medical device or its accessories before they perform sterilization, information about the correct packaging (e.g. material, composition, dimension) should be provided.

2.4.3 History of Development

For any device versions/prototypes referenced in the evidence presented in the submission, a table describing the version/name, with 4 columns (Device Name and/or Version; Description of changes from previous row; motivation for the change; list of verification/validation activities, including clinical studies, conducted using this version).

For any design verification or validation activities presented in this submission (including clinical studies) performed on any earlier versions of the subject device, include a justification for why the changes do not impact the validity of the data collected under those activities in supporting the safety and performance of the final IVD medical device design.

2.4.4   Reference and Comparison to Similar and/or Previous Generations of the Device

a) A list of the similar devices (available on local and international market) and/or previous generation of the devices (if existent) relevant to the submission. This should include any similar/previous generation devices that were previously reviewed and refused by the subject regulator.

b) Description of why they were selected.

c)  A key specification comparison, preferably in a table, between the references (similar and/or previous generation) considered and the device.

2.4.5. Substantial Equivalence Discussion

a) Identify the predicate device(s), and optionally reference devices

- trade name and model number

- Ensure the identified predicate device(s) is consistent throughout the submission

b) Include a comparison of indications for use and the technology (including features materials and principles of operation) between the predicate device(s) and subject device(s).

c) Include an analysis of why any differences between the subject device(s) and the predicate device(s) do not render the subject device(s) Not Substantially Equivalent, affect safety or effectiveness or raise different questions of safety and effectiveness.

2.5. Indications for Use and/or Intended Use and Contraindications

2.5.1. Intended Use; Intended Purpose; Intended User; Indications for Use

This section should include, as appropriate:

a) Intended Use: The statement of intended use should specify the therapeutic or diagnostic function provided by the device and may describe the medical procedure in which the device is to be used.

b) Intended Purpose: What is expected with the use of this medical device? Which results are expected?

c) Intended user and skills/knowledge/training that the user should have to operate or use the device.

d) Identify if the device is intended for single or multiple use

e) Indications for Use:

- Disease or medical condition that the device will diagnose, parameters to be monitored and other considerations related to indication for use.

- If applicable, information about patient selection criteria.

- If applicable, information about intended patient population (e.g. adults, pediatrics or newborn) or a statement that no subpopulations exist for the disease or condition for which the device is intended.

NOTES:

- The statements of intended use and purpose and the intended user and indications for use must be as presented in the labeling.

- If more than one device is included, the information should be provided for each device.

2.5.2. Intended Environment/Setting for use

- The setting where the device is intended to be used (e.g. domestic use, hospitals, medical/clinical laboratories, ambulances, medical/dental offices). Multiple options can be indicated.

- If applicable, environmental conditions that can affect the device’s safety and/or performance (e.g. temperature, humidity, power, pressure, movement).

2.5.3. Pediatric Use

- Description of any pediatric subpopulations that suffer from the disease or condition that the device is intended diagnose.

- The number of affected pediatric patients, as a whole and within each pediatric subpopulation.

OR

c) Statement that no pediatric subpopulation exists for the disease or condition for which the device is intended.

2.5.4. Contraindications for Use

If applicable, specify the disease or medical conditions that would make use of the device inadvisable due to unfavorable risk/benefit profile.

NOTE: The statement if contraindications for the device must be as presented in the labeling.

2.6. Global Market History

2.6.1. Global Market History

- Up to date indication of the markets (all countries or jurisdictions) where the device is approved for marketing, including any marketing under compassionate use regulations.

- Should include history of the marketing of the device by any other entity in as much detail as possible, acknowledging that detailed information may not be available in all cases.

- If the subject device is different in any way (e.g. design, labeling, specifications) from those approved or marketed in other jurisdiction, the differences should be described.

- The month and year of market approval in each country or jurisdiction where the device is marketed. If the device has been marketed for greater than 10 years, a statement of greater than 10 years can be made.

- For each of the markets listed in (a) above, and statement of the commercial names used in those markets OR a clear statement that the commercial names are the same in all jurisdictions.

f) State the date of data capture for the market history data.

g) If the subject device has been the subject of any previous compassionate use and/or clinical trials this should be identified and, if applicable, relevant reference numbers provided.

2.6.2. Global Incident Reports and Recalls

a)   List adverse events/incidents associated with the device and a statement of the period associated with this data.

b)   If the number of adverse events is voluminous, provide a summary by event type that state the number of reported events for each event type.

- List of the medical device recalls and/or advisory notice, and a discussion of the handling and solution given by the manufacturer in each case.

- A description of any analysis and/or corrective actions undertaken in response to items listed above.

2.6.3. Sales, Incident and Recall Rates

- A summary of the number of units sold in each country/region and a statement of the period associated with this data.

- Provide the rates calculated for each country/region, for example:

- Incident rate = # adverse events/incidents divided by # units sold x 100

- Recall rate = # recalls divided by # units sold x 100

Rates may be presented in other appropriate units such as per patient year of use or per use. In this case, methods for determining these rates should be presented and any assumptions supported.

c) Critical analyses of the rates calculated (e.g. why are they acceptable? How do they break down in terms of incidents? Is there some outlier data that has driven the rates up? Are there any trends associated with any sub-groups of the devices that are subject of the submission (e.g. size, version)?).

NOTES

i. Sales in this context should be reported as the number of units sold.

ii. The summary of sales should be broken down by components when appropriate.

2.6.4. Evaluation/Inspection Reports

Copies of Evaluation/Inspection Reports from other parties (e.g. Notified Body inspection reports).

CHAPTER 3 – ANALYTICAL PERFORMANCE AND OTHER EVIDENCES

3.1     Chapter Table of Contents

- Includes major headings for the chapter, to the level of the custom headings.

- Specifies the page number for each item referred to in the table.

3.2      Risk Management

a) A summary of the risks identified during the risk analysis process and how these risks have been controlled to an acceptable level. The summary should address

- Possible hazards for the IVD medical device for example, the risk from false positive or false negative results and the risk of delays in availability of results

- Indirect risks which may result from IVD medical device-associated hazards, for example, risk associated with instability, which could lead to erroneous results or user-related hazards, such as reagents containing infectious agents.

b) The results of the risk analysis should provide a conclusion with evidence that remaining risks are acceptable when compared to the benefits.

c)  Where a standard is followed, identify the standard

3.3    Essential Principles (EP) Checklist

a) An EP checklist established for the IVD medical devices, information about method(s) used to demonstrate conformity with each EP that applies, references for the method adopted and identification of the controlled document with evidence of conformity with each method used.

b) For the controlled documents indicated which are required for inclusion in the submission: a cross-reference of the location of such evidence within the submission.

c) If any EP indicated in the checklist does not apply to the device: a documented rationale of the non-application of each EP that does not apply.

NOTE:

Methods used to demonstrate conformity may include one or more of the following:

a) Conformity with recognized or other standards;

b) Conformity with a commonly accepted industry test method(s);

c) Conformity with an in-house test method(s);

d) The evaluation of pre-clinical and clinical evidence;

e) Comparison to a similar device already available on the market.

3.4       Standards

3.4.1         List of Standards

a) List the standards that have been complied with in full or in part in the design and manufacture of the device.

b) At a minimum should include the standard organization, standard number, standard title, year/version, and if full or partial compliance.

c)  If partial compliance, a list the sections of standard that

i. Are not applicable to the device, and/or

ii. Have been adapted, and/or

iii. Were deviated from for other reasons – discussion to accompany

3.4.2 Declaration of Conformity

The applicant is advised to prepare the Declaration of Conformity to recognized standards.

3.5 Analytical Performance

3.5.01 Stability of Specimen(s)

Information regarding and studies to support the stability, storage and where appropriate, transport, of all of the specimen type(s) identified in the labelling, including any and all recommended additives (e.g. anticoagulants) is to be provided in this section. This should include:

a) For each specimen type identified in the labelling, a description of the recommended storage parameters and when applicable, transport conditions (e.g. duration, temperatures and freeze/thaw cycles).

b) A justification on the selection of the studies performed.

c)  Provide summary of the evidence that falls within this category

d) A discussion and a conclusion to support why the evidence presented is sufficient to support the application.

OR

e)  A discussion of why this category of study is not applicable to this case.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the study results provided in this section regarding the subject device

3.5.01.1 [Study description, study identifier, date of initiation, date of completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone.

For example, the structure will look something like this

Level 3: Storage of serum samples for 7 days at 2-8ºC or 4 days at -20ºC.

Level 4: Summary

Level 4: Full Report

Level 3: Validation of 3 freeze/thaw cycles for serum samples

Level 4: Summary

Level 4: Full Report

3.5.01.1.1 Summary

A summary of the specific study described in the custom heading above.

3.5.01.1.2. Full report

The test report for the test described in the custom heading above

3.5.01.1.3 Statistical Data

This is the location for statistical data associated with the test described in the custom heading above.

3.5.02 Validation of Specimens

Studies to support the validity of specimen type(s) used in the analytical and clinical studies as representative of all of the sample type(s) identified in the labeling, including any and all recommended additives (e.g. anticoagulants),  as well as contrived specimens used in certain analytical studies are to be included in this section. This should include:

a) A list of the specimen type(s) used, including any additives (e.g. anticoagulants), in each of the analytical performance studies. If the same specimens are used for all analytical studies this can be stated and the specimen type identified.

b) For any or all of the analytical and clinical studies, if a particular specimen type(s) including additives (e.g. anticoagulants), has been chosen as representative of other specimen types identified in the labeling, this should be described and supported.

c)  If the preparation of the specimen has not followed the protocol described in the current labeling, this should be identified and validated.

d) A justification of the selection of the studies performed.

e)  Provide summary of the evidence that falls within this category

f)  A discussion and a conclusion to support why the evidence presented is sufficient to support the application

OR

g) A statement of why this category of study is not applicable to this case.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the study results provided in this section regarding the subject IVD medical device

3.5.02.1 [Study description,study identifier, date of initiation, date of completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created  for each study under the parent heading. The sub headings below would be for this study

3.5.02.1.1 Summary

A summary of the specific study described in the custom heading above.

3.5.02.1.2 Full Report

The test report for the test described in the custom heading above

3.5.02.1.3 Statistical Data

This is the location for statistical data associated with the test described in the custom heading above.

3.5.03   Metrological traceability of calibrator and control material values

Evidence that support the metrological traceability of values assigned to calibrators and trueness control materials. This should include:

a) A description of all calibrators and trueness control materials associated with the system.

b) A justification of the selection of the studies performed.

c)  Provide summary of the evidence that falls within this category, including for example, methods and acceptance criteria for the metrological traceability to reference materials and/or reference measurement procedures and a description of value assignment and validation.

d) A discussion and a conclusion to support why the evidence presented is sufficient to support the application.

OR

e)  A statement of why this category of study is not applicable to this case.

NOTES:

i. Precision control materials used during analytical studies to establish the reproducibility of a measurement procedure do not require the assessment of metrological traceability to a reference material or a reference method.

ii. The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the analytical performance study results provided in this section regarding the subject IVD medical device

3.5.03.1 [Study description, study identifier, date of initiation, date of completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone.

3.5.03.1.1Summary

A summary of the specific study described in the custom heading above.

3.5.03.1.2. Full Report

The test report for the test described in the custom heading above

3.5.03.1.3. Statistical Data

This is the location for statistical data associated with the test described in the custom heading above.

3.5.04 Accuracy of Measurement

NOTE: The general term measurement accuracy is currently used to cover both trueness and precision, whereas this term was used in the past to cover only the one component now named trueness. While measurement trueness, affected by systematic error, is normally expressed in terms of bias, measurement precision, affected by random error, is naturally expressed in terms of standard deviation. Accuracy is affected by a combination of systematic and random effects that contribute as individual components of the total error of measurement.

3.5.04.1Trueness

This section should provide a summary of information and evidence relating to the trueness of the measurement procedure. Trueness measures apply to both quantitative and qualitative assays only when a reference standard or method is available. This should include:

a) A rationale for the reference standard or method(s) used

b) A summary of the evidence that falls within this category

c)  A discussion and a conclusion to support why the evidence presented is sufficient to support the application.

OR

d) A statement of why this category of study is not applicable to this case.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the analytical performance study results provided in this section regarding the subject IVD medical device

3.5.04.1.1 [Study description, study identifier, date of initiation, date of completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone

3.5.04.1.1.1 Summary

A summary of the specific study described in the custom heading above.

3.5.04.1.1.2. Full Report

The test report for the test described in the custom heading above

3.5.04.1.1.3   Statistical Data

This is the location for statistical data associated with the test described in the custom heading above.

3.5.04.2 Precision (Repeatability and Reproducibility)

A summary of evidence that support the precision characteristics of the measurement of the subject IVD medical device is to be included in this section.

This should include:

a) A justification of the selection of the studies performed.

b) A summary of the evidence that falls within this category, including:

i. Repeatability estimates and a brief summary about the studies used to estimate, as appropriate, within-run variability.

ii. Reproducibility estimates and a brief summary of the studies used to estimate, as appropriate, variability between days, runs, sites, lots, operators (intended users) and instruments.  Such variability is also known as “Intermediate Precision”.

c) A discussion and a conclusion to support why the evidence presented is sufficient to support the application.

OR

d) A statement of why this category of study is not applicable to this case.

NOTE:

i. Studies should include the use of specimens that represent the full range of expected analyte (measured) concentrations that can be measured by the product, as claimed by the manufacturer.

ii. The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the analytical performance study results provided in this section regarding the subject IVD medical device. AT THIS LEVE

3.5.04.2.1 [Study description, study identifier, date of initiation, date of completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone

3.5.04.2.1.1 Summary

A summary of the specific study described in the custom heading above.

3.5.04.2.1.2 Full Report

The test report for the test described in the custom heading above.

3.5.04.2.1.3   Statistical Data

This is the location for statistical data associated with the test described in the custom heading above.

3.5.05       Analytical Sensitivity

Evidence that support the analytical sensitivity of the subject IVD medical device is to be included in this section. This may include studies to establish the limit of blank (LoB), limit of detection (LoD), and/or limit of quantitation (LoQ).

This should include:

a) A justification of the selection of the studies performed.

b) A summary of the evidence that falls within this category

c) A discussion and a conclusion to support why the evidence presented is sufficient to  support the application.

OR

d) A statement of why this category of study is not applicable to this case

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the analytical performance study results provided in this section regarding the subject IVD medical device

3.5.05.1 [Study description, study identifier, date of initiation, date of completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone

3.5.05.1.1 Summary

A summary of the specific study described in the custom heading above.

3.5.05.1.2 Full Report

The test report for the test described in the custom heading above.

3.5.05.1.3 Statistical Data

This is the location for statistical data associated with the test described in the custom heading above.

3.5.06 Analytic Specificity

Evidence that support the analytical specificity (interference, including as appropriate, selectivity, and cross reactivity) of the subject IVD medical device is to be included in this section. This should include:

a) A justification of the selection of the studies performed.

b) A summary of the evidence that falls within this category

c) A discussion and a conclusion to support why the evidence presented is sufficient to supporting the application or

d) A statement of why this category of study is not applicable to this case.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the analytical performance study results provided in this section regarding the subject IVD medical device

3.5.06.1Study description, study identifier, date of initiation, date of completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone.

3.5.06.1.1 Summary

A summary of the specific study described in the custom heading above.

3.5.06.1.2 Full Report

The test report for the test described in the custom heading above

3.5.06.1.3 Statistical Data

This is the location for statistical data associated with the test described in the custom heading above.

3.5.07. High Dose Hook Effect

Evidence that supports the absence of a high dose hooks effect or prozone effect. This should include:

  - A justification of the selection of the studies performed.

  - A summary of the evidence that falls within this category

  - A discussion and a conclusion to support why the evidence presented is sufficient to support the application.

OR

  - A statement of why this category of study is not applicable to this case.

NOTE:

The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the analytical performance study results provided in this section regarding the subject IVD medical device.

3.5.07.1 [Study description, study identifier, date of initiation, date of completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone.

3.5.07.1.1. Summary

A summary of the specific study described in the custom heading above.

3.5.07.1.2Full Report

The test report for the test described in the custom heading above.

3.5.071.3. Statistical Data

This is the location for statistical data associated with the test described in the custom heading above.

3.5.08. Measuring Range of the Assay

Evidence that support the measuring range (linear and non-linear measuring systems). This measuring range should include the lower limit of quantification. This should include:

a) A justification of the selection of the studies performed.

b) A summary of the evidence that falls within this category

c) A discussion and a conclusion to support why the evidence presented is sufficient to support the application.

OR

d) A statement of why this category of study is not applicable to this case.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the analytical performance study results provided in this section regarding the subject IVD medical device

3.5.08.1 [Study description, study identifier, date of initiation, date of completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created  foreach study under the parent heading. The sub headings below would be for this study alone.

3.5.08.1.1 Summary

A summary of the specific study described in the custom heading above.

3.5.08.1.2 Full Report

The test report for the test described in the custom heading above

3.5.08.1.3 Statistical Data

This is the location for statistical data associated with the test described in the custom heading above.

3.5.09    Validation of Assay Cut- off

Evidence that support the determining assay cut-off is to be included here. This should include:

a) A justification of the selection of the studies performed.

b)  A summary of the evidence that falls within this category

c)  A discussion and a conclusion to support why the evidence presented is sufficient to support the application.

OR

- A statement of why this category of study is not applicable to this case.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the analytical performance study results provided in this section regarding the subject IVD medical device.

3.5.09.1 [Study description, study identifier, date of initiation, date of completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this studyalone

3.5.09.1.1 Summary

A summary of the specific study described in the custom heading above.

3.5.09.1.2 Full Report

The test report for the test described in the custom heading above

3.5.09.1.3   Statistical Data

This is the location for statistical data associated with the test described in the custom heading above.

3.5.10         Validation of the Assay Procedure

This section should provide a summary of information and evidence supporting the validity of the assay procedure in terms of important reaction conditions (e.g. reaction time, reaction temperature, reagent volume, reading time).

This should include:

a) A justification of the selection of the studies performed.

b)  A summary of the evidence that falls within this category

c)  A discussion and a conclusion to support why the evidence presented is sufficient to support the application

OR

d) A statement of why this category of study is not applicable to this case.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the analytical performance study results provided in this section regarding the subject IVD medical device

3.5.10.1 [Study description, study identifier, date of initiation, date of completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone

3.5.10.1.1 Summary

A summary of the specific study described in the custom heading above

3.5.10.1.2 Full Report

The test report for the test described in the custom heading above.

3.5.10.1.3 Statistical Data

This is the location for statistical data associated with the test described in the custom heading above

3.6        Other Studies

3.6.1. Electrical Systems: Safety, Mechanical and Environmental Protection, and Electromagnetic Compatibility

Evidence supporting electrical safety, mechanical and environmental protection, and electromagnetic compatibility are to be included in this section.

This should include:

a) A justification of the selection of the studies performed.

b) A summary of the evidence that falls within this category

c) A discussion and a conclusion to support why the evidence presented is sufficient to support the application.

OR

d) A statement of why this category of laboratory study is not applicable to this case.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the study results provided in this section regarding the subject IVD medical device

3.6.1.1  [Study description, study identifier, date of initiation, date of Completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone.

3.6.1.1.1 Summary

A summary of the specific study described in the custom heading above

3.6.1.1.2 Full Report

The test report for the test described in the custom heading above

3.6.1.1.3 Statistical Data

3.6.2. Software/Firmware

Studies and supporting information on the software design, development process and evidence of the validation of the software, as used in the finished IVD medical device, are to be included in this section and the associated sub-sections. It should also address all of the different hardware configurations and, where applicable, operating systems identified in the labeling

3.6.2.01 Software/Firmware Description

a) Specify the name of the software

b) Specify the version of the software - The version tested must be clearly identified  and should match the release version of the software, otherwise justification must be provided.

c) Provide a description of the software including the identification of the IVD medical device features that are controlled by the software, the programming language, hardware platform, operating system (if applicable), use of Off-the-shelf software (if applicable), a description of the realization process.

d) Provide a statement about software version naming rules, specify all fields and their meanings of software version, and determine the complete version of software and its identification version used for release.

3.6.2.02 Hazard Analysis

The Hazard Analysis should take into account all device hazards associated with the IVD medical device’s intended use, including both hardware and software hazards.

NOTE:

i. This document can be in the form of an extract of the software-related items from comprehensive risk management documentation, described in ISO 14971.

ii. Hazard analysis, should address all foreseeable hazards, including those resulting from intentional or inadvertent misuse of the IVD medical device.

3.6.2.03 Software Requirement Specification

The Software Requirements Specification (SRS) documents the requirements for the software. This typically includes functional, performance, interface, design, developmental, and other requirements for the software. In effect, this document describes what the Software Device is supposed to do. For example, hardware requirements, programming language requirement, interface requirements, performance and functional requirements,

3.6.2.04 Architecture Design Chart

Detailed depiction of functional units and software modules. May include state diagrams as well as flow charts.

3.6.2.05 Software Design Specification

The Software Design Specification (SDS) describes the implementation of the requirements for the Software Device. The SDS describes how the requirements in the SRS are implemented

3.6.2.06. Traceability Analysis

A Traceability Analysis links together your product design requirements, design specifications, and testing requirements. It also provides a means of tying together identified hazards with the implementation and testing of the mitigations.

3.6.2.07 Software Life Cycle Process Description

A summary describing the software development life cycle and the processes that are in place to manage the various life cycle activities

3.6.2.08 Software Verification and Validation

a) Include an overview of all verification, validation and testing performed both inhouse and in a simulated or actual user environment prior to final release.

b) Discussion to support why the evidence presented is sufficient to support the application.

c) A statement of why this category of non-clinical laboratory study is not applicable to this case.

NOTE

i. Discussion should address all of the different hardware configurations and, where applicable, operating systems identified in the labeling.

ii. The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the non-clinical study results provided in this section regarding the subject IVD medical device

3.6.2.08.1 [Study description, study identifier, date of initiation]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone.

3.6.2.08.1.1 Summary

A summary of the specific study described in the custom heading abov

3.6.2.08.1.2 Full report

The test report for the test described in the custom heading above.

3.6.2.08.1.3 Statistical Data

3.6.2.09 Revision Level History

Revision history log, including release version number and date.

3.6.2.10Unresolved Anomalies (Bugs or Defects)

All unresolved anomalies in the release version of the software should be summarized, along with a justification for acceptability (i.e. the problem, impact on safety and performance, and any plans for correction of the problems).

3.6.2.11cybersecurity

Evidence to support the cybersecurity should be provided here. For example, but not limited to:

a) Cybersecurity vulnerabilities and risks analysis

b) Cybersecurity controls measures

c) Traceability matrix linking cybersecurity controls to the cybersecurity vulnerabilities and risks

3.6.3. Cleaning and Disinfection Validation

Contains information on the validation of cleaning and disinfection instructions for reusable devices, including evidence to support maintenance of performance when subject to this procedure over a number of cycles that is representative of the IVD medical device’s expected useful life. Information to be included in this section includes:

a) If applicable, a discussion of how the number of cycles that is representative of the IVD  medical device’s expected useful life has been determined.

b) A justification of the selection of the studies performed.

c) A summary of the evidence that falls within this category

d) A discussion and a conclusion to support why the evidence presented is sufficient to  support the application.

OR

e) A statement of why this category of laboratory study is not applicable to this case.

NOTES:

i. This applies most typically in near patient testing involving whole blood.

ii. The sponsor/applicant should explicitly address any existing regional regulatory  guidance related to the study results provided in this section regarding the subject IVD medical device.

3.6.3.1[Study description, study identifier, date of initiation]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone.

3.6.3.1.1 Summary

A summary of the specific study described in the custom heading abov

3.6.3.1.2 full report

The test report for the test described in the custom heading above.

3.6.3.1.3 Statistical Data

3.6.4 Usability/Human Factors

Studies specifically assessing the instructions and/or IVD medical device design in terms of impact of human behavior, abilities, limitations, and other characteristics on the ability of the IVD medical device to perform as intended should be included here.

This should include:

a) State the test environment and relation to the intended use environment

b) A justification of the selection of the studies performed.

c) A summary of the evidence that falls within this category

d) A discussion and conclusion to support why the evidence presented is sufficient to  support the application.

OR

e) A statement of why this category of laboratory study is not applicable to this case.

NOTES:

i. If a clinical study has been conducted that includes usability/human factors endpoints, reference to the studies and endpoints should be made, but full results do not need to be repeated and should be included in Chapter 4 – Clinical Evidence.

The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the study results provided in this section regarding the subject IVD medical device

3.6.4.1[Study description, study identifier, date of initiation]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone.

3.6.4.1.1 Summary

A summary of the specific study described in the custom heading above.

3.6.4.1.2 Full report

The test report for the test described in the custom heading above.

3.6.4.1.3 Statistical Data

3.6.5 Stability of the IVD

3.6.5.1 Claimed Shelf-life

Contains details and evidence supporting the claimed shelf-life of the IVD medical device components (e.g. reagents, calibrators/reference materials, control material, any other components susceptible to degradation)

Information provided in this section should include:

a) A description of recommended environmental conditions for storage of the IVD medical  IVD medical device (e.g. temperature, pressure, humidity, light conditions).

b) A statement of the claimed shelf-life indicated as a period of time or any other means of  appropriate quantification.

c) An indication of the packaging used in any studies conducted in support of the shelf- life. If the packaging used in the studies differs from the final device packaging, a discussion of why the evidence can be consider valid in support of the claimed shelf life.

d) A description of the simulated transport conditions that the IVD was exposed to before the start of shelf-life studies.

e) A justification of the selection of the studies performed.

f) A summary of the evidence that falls within this category

g) A discussion and a conclusion to support why the evidence presented is sufficient to  support the claimed shelf-life.

OR

h) A rationale that, for an indefinite period, the storage conditions could not affect IVD  medical device safety or performance.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the study results provided in this section regarding the subject device.

3.6.5.1[Study description, study identifier, date of initiation]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone.

3.6.5.1.1 Summary

A summary of the specific study described in the custom heading abov

3.6.5.1.2 full report

The test report for the test described in the custom heading above.

3.6.5.1.3 Statistical Data

3.6.5.2 In Use Stability

Contains details and evidence supporting the stability during actual routine use of the IVD medical device (real or simulated), including all applicable components (e.g. reagents, reaction cartridges). This may include open vial stability and/or, for automated instruments, onboard stability. Information provided in this section should include:

a) A description of recommended environmental conditions for use of the IVD medical device (e.g. temperature, pressure, humidity, light conditions).

b) A justification of the selection of the studies performed.

c) A summary of the evidence that falls within this category

d) A discussion and a conclusion to support why the evidence presented is sufficient to support the application.

OR

e) A rationale that, for an indefinite period, the storage conditions could not affect IVD medical device safety or performance.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the study results provided in this section regarding the subject IVD medical device.

3.6.5.2.1[Study description, study identifier, date of initiation]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone.

3.6.5.2.1.1 Summary

A summary of the specific study described in the custom heading abov

3.6.5.2.1.2 full report

The test report for the test described in the custom heading above.

3.6.5.2.1.3 Statistical Data

3.6.5.3 Shipping Stability

Contains details and evidence supporting the tolerance of IVD medical device, or if provided separately, the components (e.g. reagents, calibrators/reference materials) to the specified or expected shipping conditions. Information provided in this section should include:

a) An indication of environmental conditions for correct shipment of the IVD medical device (temperature, pressure, humidity, light conditions, mechanical protection etc.).

b) A justification of the selection of the studies performed.

c) A summary of the evidence that falls within this category

d) A discussion and a conclusion to support why the evidence presented is sufficient to  support the application.

OR

e) A rationale that, for an indefinite period, the storage conditions could not affect IVD  medical device safety or   performance.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the study results provided in this section regarding the subject IVD medical device.

3.6.5.3.1[Study description, study identifier, date of initiation]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone.

3.6.5.3.1.1 Summary

A summary of the specific study described in the custom heading abov

3.6.5.3.1.2 full report

The test report for the test described in the custom heading above.

3.6.5.3.1.3 Statistical Data

3.7 Analytical Performance and Other Evidence Bibliography

  - ) A listing of published studies relevant to the context of this Chapter that involve this specific IVD medical device  (e.g. analytical specificity, analytical sensitivity)

  - A legible copy of key articles, including translation where applicable to meet the regulators language requirements.

  - A discussion and a conclusion to support why the evidence presented is sufficient to support the application.

OR

  - A statement that no literature related to the IVD medical device was found.

CHAPTER 4 – CLINICAL EVIDENCE

4.1 Chapter Table of Contents

a) Includes all headings for the chapter.

b) Specifies the page number for each item referred to in the table.

4.2. Overall Clinical Evidence Summary

  - This should be a brief (1-2 page) summary of the available clinical evidence being presented in support of the submission. The document should list the evidence presented, its characteristics (e.g. well-controlled studies, partially controlled studies, studies and objective trials without matched controls, well-documented case histories conducted by qualified experts, literature review) and provide a discussion of how this is considered sufficient to support request for marketing for the requested indications. A tabular listing of clinical studies may be included in this section.

  - If any of the study IVD medical devices differ from the IVD medical devices to be marketed, including competitors’ IVD medical devices, a description of these differences and their impact on the validity of the evidence in terms of support for the application.

  - A discussion of the clinical evidence considered for the IVD medical device and support for their selection (i.e. what type of evidence was considered and why they were or were not used)

  - Discussion to support why the evidence presented is sufficient to support the application.

NOTE: Human factors testing that include patients should be included here.

4.2.1. Expected Values/Reference Ranges

This section should include information on what values to expect in healthy normal patients versus affected patients.

4.2.2 Clinical Evidence Evaluation Report

  - A clinical evidence evaluation report reviewed and signed by an expert in the relevant field that contains an objective critical evaluation of all of the clinical data submitted in relation to the IVD medical device.

  - A complete curriculum vitae, or similar documentation, to justify the manufacturer's choice of the clinical expert.

    - IVD medical Device Specific Clinical Studies

Clinical study information under this heading should be grouped by study

      - [Study description, protocol #, date of initiation, date of  completion]

This heading should be CUSTOM AND BASED ON STUDY DETAILS and created for each study under the parent heading. The sub headings below would be for this study alone.

For example, the structure will look something like this

Level 3: EU Pilot Study, CT4203, 2010-10-10

Level 4: Clinical Study Synopsis

Level 4: Clinical Study Report

Level 3: NA Controlled Study, CT4584, 2011-01-23

Level 4: Clinical Study Synopsis

Level 4: Clinical Study Report

4.2.3.1.1. Clinical Study Synopsis

a) A summary of the specific study described in the custom heading above.

b) 2-3 page summary document that presents a summary of:

i. The key characteristics of the study (e.g. title of study, investigators, sites, study period (date of enrollment/date of last completed), objectives, methods, statistical design, interpretation of design, # patients, inclusion/exclusion criteria) and

ii. Summary of the results of the analysis

iii. Summary of conclusions related to the endpoints

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the components of the clinical study synopsis.

4.2.3.1.2. Clinical Study Report

a) A clinical study report of the specific study described in the custom heading above.

NOTES:

i. The clinical study report should include elements such as the investigational plan/study protocol, protocol changes and deviations, description of patients, data quality assurance, analysis/results.

ii. The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the components of the clinical study report.

4.2.4. Clinical Literature Review and Other Reasonable Known Information

a) Clinical literature review that critically reviews available information that is published, available, or reasonably known to the applicant/sponsor that describes safety and/or performance of the IVD medical device

b) A legible copy of key articles, including translation where applicable to meet the regulators language requirements.

OR

c) A statement that no literature related to the IVD medical device was found.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to the clinical study and data provided in this section regarding the subject

IVD medical device

  - IRB Approved Informed Consent Forms

Copies of IRB approved informed consent forms are to be provided here.

  - Investigators Sites and IRB contact information

CHAPTER 5

LABELLING AND PROMOTIONAL MATERIAL

5.1. Chapter Table of Contents

a) Includes all headings for the chapter.

b) Specifies the page number for each item referred to in the table.

5.2. Product/Package Labels

Samples of the primary and secondary packaging labels but exclusive of labels for shipping.

NOTE: The sponsor/applicant should explicitly address any existing regional regulatory guidance related to labelling the subject device.

5.3. Contents of medical devices labeling

Irrespective of the class of the device, the labeling of any medical device should bear the following information:

a) Name or trade name of the device;

b) Name and complete address of the actual manufacturer of the device (street name, number, telephone, fax, e-mail, website);

c) Date of issue or latest revision of the instructions for use and, where appropriate, an identification number;

d) Sufficient details for the user to identify the device and, where these are not obvious, its intended purpose, user, and patient population of the device, and, where relevant, the contents of any packaging;

e) An indication of either the batch code/lot number (e.g., on single-use disposable devices or reagents) or model, or the serial number (e.g., on electrically-powered medical devices), where relevant, to allow appropriate actions to trace and recall the devices.

f) An unambiguous indication of the date until when the device may be used safely, expressed at least as the year and month (e.g., on devices supplied sterile, single-use disposable devices or reagents), where this is relevant. Where relevant, the storage conditions and shelf life following the first opening of the primary container, together with the storage conditions and stability of working solutions.

g) For devices other than those covered by (f) above, and as appropriate to the type of device, an indication of the dates of manufacture and expiration. This indication may be included in the batch code/lot number or serial number;

h) The information needed to verify whether the device is properly installed and can operate correctly and safely, plus details of the nature and frequency of preventative and regular maintenance and, where relevant, any quality control, replacement of consumable components, and calibration needed to ensure that the device operates properly and safely during its intended life;

i) Any warnings, precautions, limitations, or contra-indications;

j) The performance intended by the manufacturer and, where relevant, any undesirable side effects;

k) An indication on the external packaging of any special storage and /or handling conditions that apply;

l) Details of any further treatment or handling needed before the device can be used (e.g., sterilization, final assembly, calibration, preparation of reagents and/or control materials, etc.) where relevant;

m) If the device is sterile, an indication of that condition and necessary instructions in the event of damage to the sterile packaging and, where appropriate, description of methods for re-sterilization;

n) If the device has been specified by the manufacturer as intended for single-use only, an indication of that state;

o) If the device is intended for premarket clinical investigation, or for in vitro diagnostic medical devices, or for performance evaluation only, an indication of that situation;

p) If the device is intended for presentation or demonstration purposes only, an indication of that situation;

q) If the device is to be installed with or connected to other medical devices or equipment, or with dedicated software in order to operate as required for its intended use, sufficient details of its characteristics to identify the correct devices or equipment to use in order to obtain a safe combination;

r) If the device is implantable, information regarding any particular risks in connection with its implantation;

s) Information regarding the risks of reciprocal interference posed by the reasonably foreseeable presence of the device during specific investigations, evaluations, treatment, or use (e.g., electromagnetic interference from other equipment);

t) If the device is reusable, information on the appropriate processes to allow reuse, including cleaning, disinfection, packaging, and, where appropriate, the method of re-sterilization, and any restriction on the number of reuses. Where a device is supplied with the intention that it is sterilized before use, the instructions for cleaning and sterilization should be such that, if correctly followed, the device will still perform as intended by the manufacturer and comply with the Essential Principles of Safety and Performance of Medical Devices;

u) If the device emits radiation for medical purposes, details of the nature, type, and where appropriate, the intensity and distribution of this radiation,

v) Precautions and/or measures to be taken in the event of changes in the performance, or malfunction, of the device including a contact telephone number, if appropriate;

w) Precautions and/or measures to be taken as regards exposure, in reasonably foreseeable environmental conditions, to magnetic fields, external electrical influences, electrostatic discharge, pressure or variations in pressure, temperature, humidity, acceleration, thermal ignition sources, proximity to other devices, etc.;

x) If the device administers medicinal products, adequate information regarding any medicinal product(s) that the device in question is designed to administer, including any limitations in the choice of substances to be delivered;

y) Any medicinal substances or biological material incorporated into the device as an integral part of the device;

z) Any requirement for special facilities, or special training, or particular qualifications of the device user and/or third parties;

aa) Any precautions to be taken related to the disposal of the device and/or its accessories (e.g., lancets), to any consumables used with it (e.g., batteries or reagents), or to any potentially infectious substances of human or animal origin;

bb) Where relevant, for devices intended for lay persons, a statement clearly directing the user not to make any decision of medical relevance without first consulting his or her health care provider.

5.4. e-labeling

a) For eligible medical devices and stand-alone software, the applicant needs to identify which form of e-labeling is being used in case of e-labeling (e.g. electronic storage system or built-in system, website).

b) Provide details of risk management in relation to e-labeling. If this is part of the overall risk management, refer to it here

c) A description of the procedure and operations on providing IFU's when requested

d) Provide written information for user Information on webpage where IFU and further information can be found in relevant languages.

e) Description on how the requirements detailed for the website have been met.

CHAPTER 6A

QUALITY MANAGEMENT SYSTEM PROCEDURES

6A.1. Cover letter

Any modular Premarket approval submission of quality system information would need a cover letter containing the information describe in Chapter 1 under the Cover Letter heading.

6A.2. Chapter Table of Contents

a) Includes all headings for the chapter.

b) Specifies the page number for each item referred to in the table.

6A.3. Administrative

Administrative information needed to evaluate the premarket submission related to the QMS.

6A.3.1. Product Descriptive Information

Abbreviated description of the device, operating principles and overall manufacturing methods

6A.3.2. General Manufacturing Information

a) Address and contact information for all sites where the device or its components are manufactured.

b) Where applicable, addresses for all critical subcontractors, such as outsourced production, critical component or raw material production (e.g. animal tissue, drugs), and sterilization will need to be provided.

6A.4. Quality management system procedures

High level quality management system procedures for establishing and maintaining the quality management system such as the quality manual, quality policy, quality objectives, and control of documents and records ISO 13485 Elements– SOPs to satisfy clause 4.

6A.5. Management responsibilities procedures

Procedures that document the management commitment to the establishment and maintenance of the QMS by addressing quality policy, planning, responsibilities/authority/communication and management review.

ISO 13485 Elements – SOPs implementing clause 5.

6A.6. Resource management procedures

Procedures that document the adequate provision of resources to implement and maintain the QMS including human resources, infrastructure and work environment. ISO 13485 Elements – SOPs implementing clause 6.

6A.7. Product realization procedures

High level product realization procedures such as those addressing planning and customer related processes

ISO 13485 Elements – SOPs implementing sub clause 7.1 and 7.2.

6A.7.1. Design and Development Procedures

Procedures that document the systematic and controlled development of the device design from initiation of the project to transfer to production. ISO 13485 Elements – SOPs for implementing sub clause 7.3.

This section should include the following information:

a) Design Control Procedure(s) b) Design & Development Planning Procedure(s) c) Design Input Procedure(s) d) Design Output - Procedure(s) e) Design Review Procedure(s) f) Design Verification Procedure(s) g) Design Validation Procedure(s) h) Risk Analysis Procedure(s) i) Design Transfer Procedure(s) j) Design Changes Procedure(s) k) Design History File Procedure(s).

6A.7.2. Purchasing Procedures

Procedures that document that purchased products/services conform to established quality and/or product specifications.

ISO 13485 Elements – SOPs to implement sub clause 7.4.

6A.7.3. Production and service controls procedures

Procedures that document the production and service activities are carried out under controlled conditions. These SOPS address issues such as cleanliness of product and contamination control; installation and servicing activities; process validation; identification and traceability; etc. ISO 13485 Elements – SOPs implementing sub clause 7.5.

6A.7.4. Control of monitoring and measuring devices procedures

Procedure that document that monitoring and measuring equipment used in the QMS is controlled and continuously performing per the established requirements.

ISO 13485 Element- SOPs for implementing sub clause 7.6.

6A.8. QMS measurement, analysis and improvement procedures

Procedures that document how monitoring, measurement, analysis and improvement to ensure the conformity of the product and QMS, and to maintain the effectiveness of the QMS. ISO 13485 Element – SOPS for implementing clause 8.

This section should:

a) Explain how complaint handling ties to MDR procedures b) Explain how risk management is tied to the CAPA activities c) CAPA Subsystem Procedures d) Nonconforming Product Procedure(s) e) Complaint Handling Procedures f) Quality Audit Procedures.

CHAPTER 6B

QUALITY MANAGEMENT SYSTEM DEVICE SPECIFIC INFORMATION

6B.1. Chapter Table of Contents

- Includes all headings for the chapter.

- b) Specifies the page number for each item referred to in the table.

6B.2. Quality management system information

Documentation and records specific to the subject device that results from the high level quality management system procedures for establishing and maintaining the quality management system such as the quality manual, quality policy, quality objectives, and control of documents, noted in Chapter 6A.

ISO 13485 Elements – documentation specific to the subject device for the implementation of clause 4.

6B.3. Management responsibilities information

Documentation and records specific to the subject device that result from the implementation the management responsibilities procedures noted in Chapter 6A. ISO 13485 Elements – documentation specific to the subject device for the implementation of clause 5.

6B.4. Resource management information

Documentation and records specific to the subject device that result from the implementation the resource management procedures noted in Chapter 6A.

ISO 13485 Elements – documentation specific to the subject device for the implementation of clause 6.

6B.5. Product realization information

Documentation and records specific to the subject device that results from the implementation of the high level product realization procedures noted in Chapter 6A. ISO 13485 Elements – documentation specific to the subject device for the implementation of sub clause 7.1 and 7.2.

6B.5.1. Design and development information

Documentation and records specific to the subject device that results from the implementation of the design and development procedures noted in Chapter 6A. NOTE: The source of this information is the Design and Development Records (e.g. DHF - Design History File). ISO 13485 Elements – documentation specific to the subject device for the implementation of sub clause 7.3.

6B.5.2. Purchasing information

Documentation and records specific to the subject device that results from the implementation of purchasing procedures noted in Chapter 6A ISO 13485 Elements – documentation specific to the subject device for the implementation of sub clause 7.4.

List of suppliers of goods or services that affect product conformity with requirements (critical suppliers) and a description of how purchasing requirements are fulfilled for these suppliers.

6B.5.3. Production and service controls information

a) Detailed Manufacturing Flow Diagram

b) Summary of in-process acceptance activities for subject device

c) Process Validation Master Plan

d) List of processes that have not be validated

e) For each process validation considered critical to the safety and effectiveness of the device:

i. Protocols/Procedures for the validated process

ii. Process validation report

iii. The procedures for monitoring and controlling the process parameters of a validated process should be fully described.

iv. State the frequency of re-validation.

6B.5.4. Control of monitoring and measuring devices information

Documentation and records specific to the subject device that results from the implementation of the control of monitoring and measuring device procedures noted in Chapter 6A.

ISO 13485 Elements – documentation specific to the subject device for the implementation of sub clause 7.6.

6B.6. QMS measurement, analysis and improvement information

Documentation and records specific to the subject device that results from the implementation of the QMS measurement, analysis and improvement procedures noted in Chapter 6A. ISO 13485 Elements – documentation specific to the subject device for the implementation of clause 8.